## Sheet1
file ~ace/xess/voyager/v1.ecal_accmods.xs6
This spreadsheet contains the electronic cal data for V1 (FU1) located in the binder marked "FU1 Electronic Calibration" .
This spreadsheet contains some mods to clean up the original calibration data set and maybe add new points at low channels.
2026-07-29 16:47 |  | The fits to ch <512 are no longer used in the data analysis; see /home/valkyr2/voyager/gensrc/tag1_f_alldata.c (for ecal_lowz.c)
46232.6993 |  | and tag1_alldata.c (for ecal.c) Difference is in floats for ecal_lowz.c that uses randown number added between 0 and 1.
This spreadsheet also does fits to energy calibration data to compare with tag1.c and tag1_f.c.
Actual test input capacitors aren't known. Nominal capacitances are 10 pF (L1,L2,A1,A2) and 36 pF (L3,B1,B2,Cs)
Nominal conversion factors to Mev are:
10 pF | MeV = (DAC/ATTN)* |  | 0.068733
36pF | MeV = (DAC/ATTN)* |  | 0.24711
MeV number for channel 0 is extrapolated from 1st two points in calibration; also similar at 4095 in some cases
LET A L1 | [5 empty] | Need to continue substituting SLOPE and INTERCEPT functions in
21 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 12.8033 | =m | agrees with tag1_f.c
0 | - | - | -0.20539 | 1.6706 | =b
6 | 9507 | 2048 | 0.31906 | 0.33815 | 0.019081
8 | 3679 | 512 | 0.49388 | 0.49435 | 0.00047051
12 | 5972 | 512 | 0.80171 | 0.80677 | 0.0050671
16 | 8319 | 512 | 1.1168 | 1.1192 | 0.0024145
32 | 8865 | 256 | 2.3801 | 2.3689 | -0.011284
48 | 13489 | 256 | 3.6216 | 3.6185 | -0.0030999
64 | 4538 | 64 | 4.8736 | 4.8682 | -0.0053873
128 | 9201 | 64 | 9.8814 | 9.8669 | -0.014537
256 | 9246 | 32 | 19.8595 | 19.8643 | 0.0047529
512 | 4639 | 8 | 39.8565 | 39.8591 | 0.0025219
1024 | 9338 | 8 | 80.2286
2048 | 9388 | 4 | 161.3164
3072 | 14150 | 4 | 243.143
3500 | 16152 | 4 | 277.5439
4095 | - | - | 325.3675
LET A L1
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 12.699 | =m | agrees with tag1_f.c
0 | - | - | -0.15438 | 1.4343 | =b
6 | 10514 | 2048 | 0.35286 | 0.35953 | 0.0066718
8 | 3888 | 512 | 0.52194 | 0.51703 | -0.004916
12 | 6175 | 512 | 0.82896 | 0.83201 | 0.003053
16 | 8531 | 512 | 1.1452 | 1.147 | 0.0017592
24 | 13215 | 512 | 1.774 | 1.777 | 0.0029304
32 | 8976 | 256 | 2.41 | 2.4069 | -0.0030134
48 | 13677 | 256 | 3.6721 | 3.6669 | -0.0052354
64 | 4593 | 64 | 4.9327 | 4.9268 | -0.0058464
128 | 9276 | 64 | 9.962 | 9.9666 | 0.0045968
256 | 9257 | 32 | 19.8832
512 | 4637 | 8 | 39.8394
1024 | 9324 | 8 | 80.1083
2048 | 9368 | 4 | 160.9727
3072 | 14114 | 4 | 242.5244
4000 | 4603 | 1 | 316.378
4095 | 4717 | 1 | 324.2136
LET A L2
21 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.7161 | =m | agrees with tag1_f.c
0 | - | - | -0.083601 | 1.4011 | =b
2 | 1343 | 2048 | 0.045072 | 0.043665 | -0.0014078
3 | 3260 | 2048 | 0.10941 | 0.11657 | 0.0071628
4 | 5553 | 2048 | 0.18636 | 0.18948 | 0.0031146
6 | 10222 | 2048 | 0.34306 | 0.33529 | -0.0077675
8 | 3671 | 512 | 0.49281 | 0.48111 | -0.011703
12 | 5819 | 512 | 0.78117 | 0.77274 | -0.0084302
16 | 8014 | 512 | 1.0758 | 1.0644 | -0.011467
32 | 8228 | 256 | 2.2091 | 2.2309 | 0.021759
48 | 12665 | 256 | 3.4004 | 3.3974 | -0.003009
64 | 4252 | 64 | 4.5664 | 4.5639 | -0.0025387
95 | 6332 | 64 | 6.8003 | 6.824 | 0.023762
128 | 8598 | 64 | 9.2338 | 9.23 | -0.0038792
256 | 8644 | 32 | 18.5665 | 18.5621 | -0.0044123
512 | 4333 | 8 | 37.2275 | 37.2263 | -0.0011828
1024 | 8718 | 8 | 74.9018
2048 | 8756 | 4 | 150.4565
3072 | 13188 | 4 | 226.6127
4000 | 4297 | 1 | 295.3457
4095 | 4402 | 1 | 302.5627
LET A L2
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.9019 | =m | agrees with tag1_f.c
0 | - | - | -0.063665 | 0.81035 | =b
2 | 2067 | 2048 | 0.069371 | 0.085575 | 0.016204
3 | 4049 | 2048 | 0.13589 | 0.15751 | 0.021619
4 | 6413 | 2048 | 0.21523 | 0.22944 | 0.014213
6 | 11077 | 2048 | 0.37176 | 0.37331 | 0.0015501
8 | 3858 | 512 | 0.51791 | 0.51717 | -0.0007427
12 | 6007 | 512 | 0.8064 | 0.8049 | -0.0015023
16 | 8180 | 512 | 1.0981 | 1.0926 | -0.0054837
24 | 12493 | 512 | 1.6771 | 1.6681 | -0.0090165
32 | 8422 | 256 | 2.2612 | 2.2436 | -0.017651
48 | 12675 | 256 | 3.4031 | 3.3945 | -0.008607
64 | 4237 | 64 | 4.5503 | 4.5454 | -0.0049331
128 | 8526 | 64 | 9.1565 | 9.1491 | -0.0074208
256 | 8549 | 32 | 18.3625 | 18.3565 | -0.0059526
512 | 4279 | 8 | 36.7636 | 36.7713 | 0.0077235
1024 | 8601 | 8 | 73.8966
2048 | 8634 | 4 | 148.3602
3072 | 12998 | 4 | 223.3479
4000 | 4235 | 1 | 291.0843
4095 | 4338 | 1 | 298.1638
LET A L3
22 | C |  |  | ch = m*E + b |  | does not agree with tag1_f.c | From tag1_f.c | (fit should be 12 pts from ch 3-512 but can't reproduce)
Chan | DAC | ATTN | MeV | 1.9576 | =m |  | 1.9576 | =m
0 | - | - | -0.088805 | 0.26338 | =b |  | 0.26022 | =b
2 | - | - | 0.80238 | 0.88713 | 0.084743 | 0.0071813 | 0.88873 | 0.086347 | 0.0074559 | -0.0016049 | this pt scaled from LET A L3 0C assuming same slope from ch 3
3 | 10343 | 2048 | 1.248 | 1.398 | 0.14998 | 0.022495 | 1.3996 | 0.15158 | 0.022977 | -0.0016002
4 | 3822 | 512 | 1.8446 | 1.9088 | 0.064158 | 0.0041163 | 1.9104 | 0.065754 | 0.0043235 | -0.0015955
6 | 6428 | 512 | 3.1024 | 2.9305 | -0.17192 | 0.029558 | 2.9321 | -0.17034 | 0.029015 | -0.0015861
8 | 8151 | 512 | 3.934 | 3.9521 | 0.018163 | 0.0003299 | 3.9537 | 0.01974 | 0.00038966 | -0.0015766
12 | 12459 | 512 | 6.0132 | 5.9955 | -0.017697 | 0.00031317 | 5.997 | -0.016139 | 0.00026046 | -0.0015577
16 | 8317 | 256 | 8.0282 | 8.0388 | 0.010634 | 0.00011309 | 8.0404 | 0.012173 | 0.00014819 | -0.0015388
32 | 4165 | 64 | 16.0815 | 16.2122 | 0.13072 | 0.017087 | 16.2136 | 0.13218 | 0.017471 | -0.0014633
48 | 6348 | 64 | 24.5102 | 24.3855 | -0.12469 | 0.015549 | 24.3869 | -0.12331 | 0.015205 | -0.0013877
64 | 8465 | 64 | 32.6842 | 32.5589 | -0.12527 | 0.015693 | 32.5602 | -0.12396 | 0.015366 | -0.0013122
128 | 8459 | 32 | 65.322 | 65.2523 | -0.069669 | 0.0048538 | 65.2533 | -0.068659 | 0.0047141 | -0.00101
256 | 4223 | 8 | 130.4432 | 130.6392 | 0.19598 | 0.038409 | 130.6396 | 0.19639 | 0.038568 | -0.00040561
512 | 8465 | 8 | 261.4733 | 261.4129 | -0.060379 | 0.0036456 | 261.4121 | -0.061182 | 0.0037432 | 0.00080316
1024 | 8485 | 4 | 524.1821 |  |  | 0.15216 |  |  | 0.15218
2048 | 4251 | 1 | 1050.4646
2500 | 5198 | 1 | 1284.4778
2600 | 5406 | 1 | 1335.8767
3073 | 6403 | 1 | 1582.2453
4000 | 8365 | 1 | 2067.0752
4090 | 8555 | 1 | 2114.026
4095 | 8568 | 1 | 2117.2385
LET A L3
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.9941 | =m | agrees with tag1_f.c
0 | - | - | 0.1127 | -0.22587 | =b
2 | 8320 | 2048 | 1.0039 | 1.1162 | 0.11236
3 | 12013 | 2048 | 1.4495 | 1.6177 | 0.16825
4 | 4213 | 512 | 2.0333 | 2.1192 | 0.085871
6 | 6786 | 512 | 3.2752 | 3.1222 | -0.15298
8 | 8513 | 512 | 4.1087 | 4.1252 | 0.016479
12 | 12848 | 512 | 6.2009 | 6.1311 | -0.069805
16 | 8485 | 256 | 8.1903 | 8.1371 | -0.053287
24 | 12610 | 256 | 12.1721 | 12.149 | -0.023147
32 | 4181 | 64 | 16.1432 | 16.1608 | 0.017611
48 | 6275 | 64 | 24.2284 | 24.1846 | -0.043733
64 | 8351 | 64 | 32.244 | 32.2084 | -0.035578
128 | 8346 | 32 | 64.4494 | 64.3036 | -0.14582
256 | 4155 | 8 | 128.3428 | 128.4938 | 0.15109
512 | 8317 | 8 | 256.9017 | 256.8744 | -0.027317
1024 | 8328 | 4 | 514.483
2048 | 4166 | 1 | 1029.4603
3072 | 6274 | 1 | 1550.3681
4000 | 8197 | 1 | 2025.5607
4095 | 8394 | 1 | 2074.2413
LET B L1
22 | C |  |  | ch = m*E + b |  | agrees with tag1_f.c
Chan | DAC | ATTN | MeV | 12.7056 | =m
0 | - | - | -0.25063 | 2.0327 | =b
6 | 8717 | 2048 | 0.29255 | 0.31225 | 0.019694
8 | 3528 | 512 | 0.47361 | 0.46966 | -0.0039567
12 | 5861 | 512 | 0.7868 | 0.78448 | -0.0023256
32 | 8796 | 256 | 2.3616 | 2.3586 | -0.0030303 | pt at ch 16 omitted from tag1_f.c; not in line with others
48 | 13508 | 256 | 3.6267 | 3.6179 | -0.0088563
64 | 4550 | 64 | 4.8865 | 4.8772 | -0.0093127
100 | 7182 | 64 | 7.7131 | 7.7106 | -0.0025532 | this pt taken at 27C
128 | 9222 | 64 | 9.904 | 9.9143 | 0.010341
256 | 9266 | 32 | 19.9025
512 | 4644 | 8 | 39.8995
1024 | 9343 | 8 | 80.2716
2048 | 9385 | 4 | 161.2648
4000 | 4612 | 1 | 316.9966
4095 | 4725 | 1 | 324.7634
LET B L1
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 12.6352 | =m | agrees with tag1_f.c
0 | - | - | -0.17599 | 1.6168 | =b
6 | 10143 | 2048 | 0.34041 | 0.34691 | 0.0064958
8 | 3818 | 512 | 0.51254 | 0.50519 | -0.0073512
12 | 6117 | 512 | 0.82117 | 0.82177 | 0.00059662
16 | 8447 | 512 | 1.134 | 1.1383 | 0.0043829
24 | 13184 | 512 | 1.7699 | 1.7715 | 0.0016186
32 | 8956 | 256 | 2.4046 | 2.4046 | 6.2578e-05
48 | 13694 | 256 | 3.6767 | 3.6709 | -0.0057344
64 | 4600 | 64 | 4.9402 | 4.9372 | -0.0029398
128 | 9311 | 64 | 9.9996 | 10.0024 | 0.0028689
256 | 9315 | 32 | 20.0077
512 | 4659 | 8 | 40.0284
1024 | 9358 | 8 | 80.4004
2048 | 9395 | 4 | 161.4366
3072 | 14148 | 4 | 243.1086
4000 | 4612 | 1 | 316.9966
4095 | - | - | 324.5606
LET B L2
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.5904 | =m | agrees with tag1_f.c
0 | - | - | -0.0885 | 1.4083 | =b
2 | 1271 | 2048 | 0.042656 | 0.043537 | 0.00088092 | omitted pt at ch 1 with 0 MeV as did tag1_f.c
3 | 3225 | 2048 | 0.10823 | 0.11712 | 0.0088838
4 | 5447 | 2048 | 0.18281 | 0.1907 | 0.0078922
6 | 10007 | 2048 | 0.33585 | 0.33786 | 0.0020161
8 | 3647 | 512 | 0.48959 | 0.48502 | -0.0045648
12 | 5834 | 512 | 0.78318 | 0.77935 | -0.0038324
16 | 8027 | 512 | 1.0776 | 1.0737 | -0.0039054
32 | 8434 | 256 | 2.2644 | 2.251 | -0.01346
48 | 12757 | 256 | 3.4251 | 3.4283 | 0.0031625
64 | 4289 | 64 | 4.6062 | 4.6056 | -0.00061994
128 | 8672 | 64 | 9.3133 | 9.3148 | 0.0014337
256 | 8720 | 32 | 18.7297 | 18.7331 | 0.0033932
512 | 4373 | 8 | 37.5712 | 37.5699 | -0.0012796
1024 | 8798 | 8 | 75.5891
2048 | 8838 | 4 | 151.8656
3072 | 13309 | 4 | 228.6919
4000 | 4338 | 1 | 298.1638
4095 | 4443 | 1 | 305.3807
LET B L2
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.7667 | =m | agrees with tag1_f.c
0 | - | - | -0.059806 | 0.82605 | =b
2 | 2052 | 2048 | 0.068867 | 0.085274 | 0.016407
3 | 3969 | 2048 | 0.1332 | 0.15791 | 0.024709
4 | 6370 | 2048 | 0.21378 | 0.23055 | 0.016768
6 | 11022 | 2048 | 0.36991 | 0.37583 | 0.0059199
8 | 3860 | 512 | 0.51818 | 0.52111 | 0.0029249
12 | 6040 | 512 | 0.81083 | 0.81166 | 0.00082794
16 | 8222 | 512 | 1.1038 | 1.1022 | -0.0015375
24 | 12622 | 512 | 1.6944 | 1.6833 | -0.011101
32 | 8518 | 256 | 2.287 | 2.2644 | -0.022544
48 | 12831 | 256 | 3.445 | 3.4267 | -0.018313
64 | 4288 | 64 | 4.6051 | 4.5889 | -0.01623
128 | 8605 | 64 | 9.2414 | 9.2378 | -0.0036004
256 | 8630 | 32 | 18.5364 | 18.5355 | -0.00089512
512 | 4321 | 8 | 37.1244 | 37.1311 | 0.0066634
1024 | 8685 | 8 | 74.6183
2048 | 8719 | 4 | 149.8208
3072 | 13125 | 4 | 225.5302
4000 | 4276 | 1 | 293.9023
4095 | 4380 | 1 | 301.0505
LET B L3
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.8545 | =m | agrees with tag1_f.c
0 | - | - | 0.057434 | 0.30118 | =b
2 | - | - | 0.94862 | 0.91604 | -0.032586 | this pt scaled from LA3 room assuming same slope
3 | - | - | 1.3942 | 1.4553 | 0.06104 | this pt scaled from LA3 room assuming same slope
4 | - | - | 1.9909 | 1.9945 | 0.0036007 | this pt scaled from LA3 room assuming same slope
6 | 6731 | 512 | 3.2486 | 3.0729 | -0.17571
8 | 8499 | 512 | 4.1019 | 4.1514 | 0.049429
12 | 13047 | 512 | 6.297 | 6.3082 | 0.011278
16 | 8733 | 256 | 8.4297 | 8.4651 | 0.035387
32 | 4388 | 64 | 16.9425 | 17.0926 | 0.15017
48 | 6689 | 64 | 25.8269 | 25.7202 | -0.10669
64 | 8917 | 64 | 34.4294 | 34.3477 | -0.08168
128 | 8919 | 32 | 68.8742 | 68.8578 | -0.0164
256 | 4456 | 8 | 137.6403 | 137.878 | 0.23771
512 | 8936 | 8 | 276.0219 | 275.9184 | -0.1035
1024 | 8962 | 4 | 553.65
2048 | 4490 | 1 | 1109.5239
3072 | 6766 | 1 | 1671.9463 | is 1671.845801 in tag1_f.c
4000 | 8844 | 1 | 2185.4408
4095 | 9059 | 1 | 2238.5695
LET B L3
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.8903 | =m | agrees with tag1_f.c
0 | - | - | 0.26135 | -0.3917 | =b
2 | - | - | 1.1525 | 1.2653 | 0.11274 | this pt scaled from LA3 cold assumig same slope
3 | - | - | 1.5981 | 1.7943 | 0.19617 | this pt scaled from LA3 cold assumig same slope
4 | 4521 | 512 | 2.182 | 2.3233 | 0.14133
6 | 7186 | 512 | 3.4682 | 3.3814 | -0.086836
8 | 8994 | 512 | 4.3408 | 4.4394 | 0.098615
12 | 13639 | 512 | 6.5827 | 6.5556 | -0.027117
16 | 9040 | 256 | 8.7261 | 8.6717 | -0.054391
24 | 13346 | 256 | 12.8825 | 12.9039 | 0.021373
32 | 4430 | 64 | 17.1046 | 17.1361 | 0.031498
48 | 6636 | 64 | 25.6222 | 25.6006 | -0.021612
64 | 8856 | 64 | 34.1938 | 34.0651 | -0.12878
128 | 8808 | 32 | 68.017 | 67.9229 | -0.094109
256 | 4386 | 8 | 135.4781 | 135.6386 | 0.16056
512 | 8777 | 8 | 271.1106 | 271.07 | -0.040539
1024 | 8789 | 4 | 542.9624
2048 | 4403 | 1 | 1088.0253
3072 | 6627 | 1 | 1637.598
4000 | 8659 | 1 | 2139.7255
4095 | 8868 | 1 | 2191.3715
LET C L1
22 | C |  |  | ch = m*E + b |  | agrees with tag1_f.c but fit is to ch32-512
Chan | DAC | ATTN | MeV | 13.995 | =m
0 | - | - | -0.1574 | -0.18039 | =b
4 | 5452 | 2048 | 0.18297 | 0.29871 | 0.11573
6 | 10523 | 2048 | 0.35316 | 0.44161 | 0.088451
8 | 3868 | 512 | 0.51926 | 0.58452 | 0.065265
12 | 5956 | 512 | 0.79956 | 0.87034 | 0.07078
16 | 8282 | 512 | 1.1118 | 1.1562 | 0.044344
32 | 8483 | 256 | 2.2776 | 2.2994 | 0.021832
48 | 12773 | 256 | 3.4294 | 3.4427 | 0.013281
64 | 4268 | 64 | 4.5836 | 4.5859 | 0.0023146
128 | 8561 | 64 | 9.1941 | 9.159 | -0.035109
256 | 8530 | 32 | 18.3216 | 18.3051 | -0.016523
512 | 4258 | 8 | 36.5831 | 36.5973 | 0.014205
1024 | 8547 | 8 | 73.4326
2048 | 8575 | 4 | 147.3464
3072 | 12900 | 4 | 221.6639
3500 | 14718 | 4 | 252.9031 | omitted pt at ch 4000, as did tag1_f.c
4095 | - | - | 296.3313
LET C L1
0 | C |  |  | ch = m*E + b |  | agrees with tag1_f.c except see errors in tag1_f.c from ch256-4095
Chan | DAC | ATTN | MeV | 13.7763 | =m
0 | - | - | -0.10293 | 0.39879 | =b
4 | - | - | 0.23744 | 0.26141 | 0.023963 | this pt scaled from LET C L1 22 C assuming same slope
6 | 12146 | 2048 | 0.40763 | 0.40658 | -0.0010479
8 | 4158 | 512 | 0.55819 | 0.55176 | -0.0064255
12 | 6234 | 512 | 0.83688 | 0.84212 | 0.0052382
16 | 8438 | 512 | 1.1328 | 1.1325 | -0.00028144
24 | 12745 | 512 | 1.7109 | 1.7132 | 0.002238
32 | 8571 | 256 | 2.3012 | 2.2939 | -0.0073245
48 | 12920 | 256 | 3.4689 | 3.4553 | -0.013562
64 | 4315 | 64 | 4.6341 | 4.6167 | -0.017383
128 | 8611 | 64 | 9.2478 | 9.2624 | 0.014586
256 | 8571 | 32 | 18.4097 | is 18.321625 in tag1_f.c from 22 C cal
512 | 4280 | 8 | 36.7722 | is 36.583099 in tag1_f.c from 22 C cal
1024 | 8583 | 8 | 73.7419 | is 73.432587 in tag1_f.c from 22 C cal
2048 | 8607 | 4 | 147.8962 | is 147.346313 in tag1_f.c from 22C cal
3072 | 12948 | 4 | 222.4887 | is 221.663818 in tag1_f.c from 22 C cal
4000 | 4217 | 1 | 289.8471 | not in tag1_f.c; 3500 = 252.902954 is in tag1_f.c from 22 C cal
4095 | 4319 | 1 | 296.8578 | is 296.331192 in tag1_f.c from 22C cal
LET C L2
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.141 | =m | agrees with tag1_f.c
0 | - | - | -0.039199 | 0.8699 | =b
1 | 573 | 2048 | 0.01923 | 0.0099001 | -0.0093304
2 | 2314 | 2048 | 0.07766 | 0.085998 | 0.0083373
3 | 3852 | 2048 | 0.12928 | 0.16209 | 0.032818
4 | 6385 | 2048 | 0.21429 | 0.23819 | 0.023905
6 | 11167 | 2048 | 0.37478 | 0.39039 | 0.015611
8 | 3979 | 512 | 0.53416 | 0.54258 | 0.0084247
16 | 8509 | 512 | 1.1423 | 1.1514 | 0.0090782
32 | 8911 | 256 | 2.3925 | 2.3689 | -0.023579
48 | 13447 | 256 | 3.6104 | 3.5865 | -0.023882
64 | 4496 | 64 | 4.8285 | 4.804 | -0.024455
128 | 9027 | 64 | 9.6946 | 9.6743 | -0.020301
256 | 9046 | 32 | 19.43 | 19.4147 | -0.015214
512 | 4525 | 8 | 38.8771 | 38.8957 | 0.018586
1024 | 9084 | 8 | 78.0463
2048 | 9114 | 4 | 156.6081
3072 | 13725 | 4 | 235.8401
4000 | 4472 | 1 | 307.374
4095 | 4581 | 1 | 314.8659
LET C L2
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.247 | =m | agrees with tag1_f.c
0 | - | - | -0.064538 | 0.63824 | =b
2 | 2493 | 2048 | 0.083668 | 0.1028 | 0.01913
3 | 4701 | 2048 | 0.15777 | 0.17829 | 0.020515
4 | 7337 | 2048 | 0.24624 | 0.25377 | 0.007537
6 | 12094 | 2048 | 0.40589 | 0.40475 | -0.0011357
8 | 4152 | 512 | 0.55738 | 0.55573 | -0.0016532
12 | 6468 | 512 | 0.86829 | 0.85768 | -0.010608
16 | 8719 | 512 | 1.1705 | 1.1596 | -0.010838
24 | 13254 | 512 | 1.7793 | 1.7635 | -0.015727
32 | 8900 | 256 | 2.3895 | 2.3675 | -0.022092
48 | 13329 | 256 | 3.5787 | 3.5753 | -0.0034101
64 | 4441 | 64 | 4.7694 | 4.7831 | 0.013661
128 | 8948 | 64 | 9.6097 | 9.6144 | 0.004621
256 | 8940 | 32 | 19.2023 | 19.2769 | 0.074605
512 | 4468 | 8 | 38.3874
1024 | 8964 | 8 | 77.0153
2048 | 8991 | 4 | 154.4946
2500 | 10993 | 4 | 188.8955
3072 | 13534 | 4 | 232.5581
4000 | 4408 | 1 | 302.9751
4095 | 4515 | 1 | 310.3295
LET C L3
22 | C |  |  | ch = m*E + b |  |  |  | Interpolate low channels
Chan | DAC | ATTN | MeV | 1.7969 | =m | agrees with tag1_f.c |  | Chan | MeV
0 | - | - | -1.338 | 0.86677 | =b |  |  | 1 | -0.62607 | 0.71192
2 | - | - | 0.11619 | 0.63067 | 0.51447 | this pt scaled from LET C L3 0C assuming same slope |  | 4 | 1.5097 | 0
3 | - | - | 0.84329 | 1.1872 | 0.3439 | this pt scaled from LET C L3 0C assuming same slope |  | 4 | 1.5097 | -1.4238
4 | 3128 | 512 | 1.5097 | 1.7437 | 0.23402 |  |  | 2 | 0.085849 | 0.71192
6 | 5486 | 512 | 2.6477 | 2.8567 | 0.209 |  |  | 3 | 0.79777 | 0.71192
8 | 7847 | 512 | 3.7873 | 3.9698 | 0.18254 |  |  | 4 | 1.5097 | 0.56903
16 | 8562 | 256 | 8.2647 | 8.4219 | 0.15728 |  |  | 5 | 2.0787 | 0.56903
24 | 13585 | 256 | 13.1132 | 12.8741 | -0.23913 |  |  | 6 | 2.6477 | 0.56975
32 | 4537 | 64 | 17.5178 | 17.3263 | -0.19151 |  |  | 7 | 3.2175 | 0.56975
48 | 6845 | 64 | 26.4292 | 26.2306 | -0.19859 |  |  | 8 | 3.7873 | 0.55968
64 | 9150 | 64 | 35.329 | 35.1349 | -0.19409 |  |  | 9 | 4.3469 | 0.55968
128 | 9180 | 32 | 70.8897 | 70.7522 | -0.13748 |  |  | 10 | 4.9066 | 0.55968
256 | 4591 | 8 | 141.8103 | 141.9868 | 0.17653 |  |  | 11 | 5.4663 | 0.55968
512 | 9209 | 8 | 284.4545 | 284.4559 | 0.0014423 |  |  | 12 | 6.026 | 0.55968
1024 | 9241 | 4 | 570.8859 |  |  |  |  | 13 | 6.5856 | 0.55968
2048 | 4636 | 1 | 1145.602 |  |  |  |  | 14 | 7.1453 | 0.55968
2200 | 4985 | 1 | 1231.8434 |  |  |  |  | 15 | 7.705 | 0.55968
2400 | 5442 | 1 | 1344.7726 |  |  |  |  | 16 | 8.2647 | 0.60607
2600 | 5903 | 1 | 1458.6903 |  |  |  |  | 17 | 8.8707 | 0.60607
3072 | 6989 | 1 | 1727.0518 |  |  |  |  | 18 | 9.4768 | 0.60607
4000 | 9143 | 1 | 2259.3267 |  |  |  |  | 19 | 10.0829
4090 | 9355 | 1 | 2311.7141
4095 | 9367 | 1 | 2314.6794
LET C L3
0 | C |  |  | ch = m*E + b |  | agrees with tag1_f.c but fit is to ch 6-512 |  | Interpolate low channels |  | Interp to temp | times gain
Chan | DAC | ATTN | MeV | 1.833 | =m |  |  | Chan | MeV | 6.28 | 0.9035
0 | - | - | -0.9958 | 0.25715 | =b |  |  | 1 | -0.26871 | -0.37072 | -0.33494 | 0.65301
2 | 3799 | 2048 | 0.45838 | 0.95084 | 0.49245 |  |  | 2 | 0.45838 | 0.35204 | 0.31807 | 0.65301
3 | 9825 | 2048 | 1.1855 | 1.4964 | 0.31092 |  |  | 3 | 1.1855 | 1.0748 | 0.97108 | 0.61383
4 | 3837 | 512 | 1.8519 | 2.042 | 0.19008 |  |  | 4 | 1.8519 | 1.7542 | 1.5849 | 0.58921
6 | 6677 | 512 | 3.2226 | 3.1331 | -0.089482 |  |  | 5 | 2.5372 | 2.4063 | 2.1741 | 0.58921
8 | 8653 | 512 | 4.1763 | 4.2242 | 0.047952 |  |  | 6 | 3.2226 | 3.0585 | 2.7633 | 0.45479
12 | 13301 | 512 | 6.4196 | 6.4065 | -0.013095 |  |  | 7 | 3.6994 | 3.5618 | 3.2181 | 0.45479
16 | 8887 | 256 | 8.5784 | 8.5887 | 0.010319 |  |  | 8 | 4.1763 | 4.0652 | 3.6729 | 0.50641
24 | 13346 | 256 | 12.8825 | 12.9532 | 0.070662 |  |  | 9 | 4.7371 | 4.6257 | 4.1793 | 0.50641
32 | 4461 | 64 | 17.2243 | 17.3177 | 0.093358 |  |  | 10 | 5.2979 | 5.1862 | 4.6857 | 0.50641
48 | 6750 | 64 | 26.0624 | 26.0467 | -0.015692 |  |  | 11 | 5.8587 | 5.7467 | 5.1921 | 0.50641
64 | 9039 | 64 | 34.9004 | 34.7757 | -0.12474 |  |  | 12 | 6.4196 | 6.3072 | 5.6986 | 0.49278
128 | 9036 | 32 | 69.7777 | 69.6917 | -0.08603 |  |  | 13 | 6.9593 | 6.8526 | 6.1913 | 0.49278
256 | 4512 | 8 | 139.37 | 139.5236 | 0.15356 |  |  | 14 | 7.499 | 7.398 | 6.6841 | 0.49278
512 | 9040 | 8 | 279.2343 | 279.1875 | -0.04681 |  |  | 15 | 8.0387 | 7.9434 | 7.1769 | 0.49278
1024 | 9064 | 4 | 559.9513 |  |  |  |  | 16 | 8.5784 | 8.4888 | 7.6697 | 0.50365
2048 | 4546 | 1 | 1123.3621 |  |  |  |  | 17 | 9.1164 | 9.0463 | 8.1733 | 0.50365
3072 | 6849 | 1 | 1692.4564 |  |  |  |  | 18 | 9.6544 | 9.6037 | 8.677
4000 | 8954 | 1 | 2212.6229 |  |  |  |  | 19 | 10.1924 | 10.1612 | 9.1806
4095 | 9170 | 1 | 2265.9987
LET D L1
22 | C |  |  | ch = m*E + b |  | agrees with tag1_f.c but fit is to ch 24-512
Chan | DAC | ATTN | MeV | 13.6315 | =m
0 | - | - | -0.16291 | 0.20275 | =b
4 | - | - | 0.17747 | 0.27856 | 0.10109 | this pt scaled from LET C L1 22C assuming same slope
6 | 10359 | 2048 | 0.34766 | 0.42528 | 0.077625
8 | 3666 | 512 | 0.49214 | 0.572 | 0.079864
16 | 8079 | 512 | 1.0846 | 1.1589 | 0.074321
24 | 12986 | 512 | 1.7433 | 1.7458 | 0.0024616
32 | 8710 | 256 | 2.3385 | 2.3326 | -0.0059003
48 | 13064 | 256 | 3.5075 | 3.5064 | -0.0011451
64 | 4355 | 64 | 4.6771 | 4.6801 | 0.0030732
128 | 8727 | 64 | 9.3724 | 9.3752 | 0.0027629
256 | 8737 | 32 | 18.7663 | 18.7652 | -0.0010795
512 | 4370 | 8 | 37.5454 | 37.5452 | -0.00017276
728 | 6227 | 8 | 53.5
1024 | 8781 | 8 | 75.4431 | is 74.443024 in tag1_f.c
2048 | 8818 | 4 | 151.5219
3072 | 13275 | 4 | 228.1076
4000 | 4327 | 1 | 297.4077
4090 | 4426 | 1 | 304.2123
4095 | 4431 | 1 | 304.5559
LET D L1
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 13.8 | =m | agrees with tag1_f.c
0 | - | - | -0.1032 | 0.31209 | =b
4 | - | - | 0.23718 | 0.26724 | 0.030064 | 0.00090385 | this pt scaled from LET C L1 22C assuming same slope
6 | 12138 | 2048 | 0.40736 | 0.41217 | 0.0048036 | 2.3075e-05
8 | 4103 | 512 | 0.5508 | 0.55709 | 0.0062913 | 3.958e-05
12 | 6307 | 512 | 0.84668 | 0.84695 | 0.00027218 | 7.4079e-08
16 | 8498 | 512 | 1.1408 | 1.1368 | -0.0040018 | 1.6014e-05
24 | 12829 | 512 | 1.7222 | 1.7165 | -0.0057032 | 3.2526e-05
32 | 8587 | 256 | 2.3055 | 2.2962 | -0.009284 | 8.6192e-05
48 | 12837 | 256 | 3.4466 | 3.4556 | 0.0090608 | 8.2098e-05
64 | 4286 | 64 | 4.603 | 4.6151 | 0.012102 | 0.00014645
128 | 8624 | 64 | 9.2618 | 9.2527 | -0.0090272 | 8.149e-05
256 | 8631 | 32 | 18.5386 | 18.5281 | -0.010475 | 0.00010972
512 | 4315 | 8 | 37.0729 | 37.0788 | 0.0059613 | 3.5537e-05
1024 | 8666 | 8 | 74.455
2048 | 8699 | 4 | 149.4771
3072 | 13093 | 4 | 224.9803
4000 | 4266 | 1 | 293.215
4095 | 4369 | 1 | 300.2945
LET D L2
22 | C |  |  | ch = m*E + b |  | agrees with tag1_f.c but fit is to ch 24-512
Chan | DAC | ATTN | MeV | 12.7511 | =m
0 | - | - | -0.053664 | 0.2502 | =b
1 | - | - | 0.0047657 | 0.058803 | 0.054037 | this pt scaled from LET C L2 22C assuming same slope
2 | - | - | 0.063195 | 0.13723 | 0.074032 | this pt scaled from LET C L2 22C assuming same slope
3 | - | - | 0.11481 | 0.21565 | 0.10084 | this pt scaled from LET C L2 22C assuming same slope
4 | - | - | 0.19982 | 0.29408 | 0.094255 | this pt scaled from LET C L2 22C assuming same slope
6 | 10736 | 2048 | 0.36031 | 0.45093 | 0.090616
8 | 3861 | 512 | 0.51832 | 0.60778 | 0.08946
16 | 8582 | 512 | 1.1521 | 1.2352 | 0.083092 |  | 0.018488
24 | 13795 | 512 | 1.8519 | 1.8626 | 0.010676 |  | 0.074663
32 | 9261 | 256 | 2.4865 | 2.49 | 0.0035031 |  | 0.12429
48 | 13938 | 256 | 3.7422 | 3.7448 | 0.0025809 |  | 0.20602
64 | 4654 | 64 | 4.9982 | 4.9996 | 0.0013901
128 | 9338 | 64 | 10.0286 | 10.0188 | -0.0098165
256 | 9347 | 32 | 20.0765 | 20.0571 | -0.019342
512 | 4670 | 8 | 40.1229 | 40.1339 | 0.011008
1024 | 9367 | 8 | 80.4778
2048 | 9389 | 4 | 161.3335
3072 | 14135 | 4 | 242.8852
4000 | 4604 | 1 | 316.4467
4095 | 4716 | 1 | 324.1448
LET D L2
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 12.8567 | =m | agrees with tag1_f.c
0 | - | - | -0.061148 | 0.65301 | =b
2 | 2536 | 2048 | 0.085111 | 0.10477 | 0.019659 | 0.00038647
3 | 4715 | 2048 | 0.15824 | 0.18255 | 0.02431 | 0.00059097
4 | 7537 | 2048 | 0.25295 | 0.26033 | 0.007381 | 5.4479e-05
6 | 12280 | 2048 | 0.41213 | 0.41589 | 0.0037619 | 1.4152e-05
8 | 4261 | 512 | 0.57201 | 0.57145 | -0.00056203 | 3.1588e-07
12 | 6636 | 512 | 0.89084 | 0.88257 | -0.0082701 | 6.8395e-05
16 | 8970 | 512 | 1.2042 | 1.1937 | -0.010474 | 0.00010971
24 | 13646 | 512 | 1.8319 | 1.8159 | -0.015956 | 0.0002546
32 | 9176 | 256 | 2.4636 | 2.4382 | -0.025466 | 0.00064851
48 | 13751 | 256 | 3.692 | 3.6827 | -0.0093127 | 8.6727e-05
64 | 4586 | 64 | 4.9251 | 4.9272 | 0.0020076 | 4.0303e-06
128 | 9211 | 64 | 9.8922 | 9.9051 | 0.012922 | 0.00016698
256 | 9222 | 32 | 19.808
512 | 4605 | 8 | 39.5644
1024 | 9233 | 8 | 79.3265
2048 | 9255 | 4 | 159.031
3072 | 13922 | 4 | 239.2252
4000 | 4534 | 1 | 311.6354
4095 | 4644 | 1 | 319.1961
LET D L3
22 | C |  |  | ch = m*E + b |  |  |  | Interpolate low channels
Chan | DAC | ATTN | MeV | 1.9192 | =m | agrees with tag1_f.c |  | Chan | MeV
0 | - | - | -0.94259 | 0.77347 | =b |  |  | 1 | -0.35788 | 0.58471
2 | - | - | 0.28283 | 0.6391 | 0.35627 | 0.12693 | scaled from LET D L3 0C assuming same slope | 2 | 0.22684 | 0.58471
3 | - | - | 0.89553 | 1.1602 | 0.26463 | 0.070027 | scaled from LET D L3 0C assuming same slope | 3 | 0.81155 | -0.58471
4 | 2893 | 512 | 1.3963 | 1.6812 | 0.28495 | 0.081197 |  | 2 | 0.22684 | 0.58471
6 | 5385 | 512 | 2.599 | 2.7233 | 0.12434 | 0.015461 |  | 3 | 0.81155 | 0.58471
8 | 7565 | 512 | 3.6511 | 3.7655 | 0.11432 | 0.013068 |  | 4 | 1.3963 | 0.60137
16 | 8140 | 256 | 7.8573 | 7.934 | 0.076627 | 0.0058716 |  | 5 | 1.9976 | 0.60137
24 | 12789 | 256 | 12.3449 | 12.1024 | -0.24244 | 0.058778 |  | 6 | 2.599 | 0.52607
32 | 4282 | 64 | 16.5332 | 16.2709 | -0.26227 | 0.068787 |  | 7 | 3.1251 | 0.52607
48 | 6426 | 64 | 24.8114 | 24.6079 | -0.20348 | 0.041405 |  | 8 | 3.6511 | 0.52577
64 | 8560 | 64 | 33.051 | 32.9449 | -0.10608 | 0.011252 |  | 9 | 4.1769 | 0.52577
128 | 8557 | 32 | 66.0788 | 66.2928 | 0.21404 | 0.045812 |  | 10 | 4.7027 | 0.52577
256 | 4275 | 8 | 132.0494 |  |  |  |  | 11 | 5.2285 | 0.52577
512 | 8569 | 8 | 264.6857 |  |  |  |  | 12 | 5.7542 | 0.52577
1024 | 8592 | 4 | 530.7923 |  |  |  |  | 13 | 6.28 | 0.52577
2048 | 4308 | 1 | 1064.5499 |  |  |  |  | 14 | 6.8058 | 0.52577
3072 | 6490 | 1 | 1603.7439 |  |  |  |  | 15 | 7.3316 | 0.52577
4000 | 8490 | 1 | 2097.9639 |  |  |  |  | 16 | 7.8573 | 0.56094
4095 | 8695 | 1 | 2148.6215 |  |  |  |  | 17 | 8.4183 | 0.56094
[8 empty] | 18 | 8.9792 | -8.9792
LET D L3
0 | C |  |  | ch = m*E + b |  |  |  | Interpolate low channels |  | Interp to temp | times gain
Chan | DAC | ATTN | MeV | 1.9563 | =m | agrees with tag1_f.c |  | Chan | MeV | 6.28 | 0.9695
0 | - | - | -0.68197 | 0.31533 | =b |  | -0.93004 | 1 | -0.069258 | -0.15165 | -0.14702 | 0.58627
2 | 4504 | 2048 | 0.54345 | 0.86117 | 0.31772 | 0.10095 | 0.36343 | 2 | 0.54345 | 0.45307 | 0.43925 | 0.58627
3 | 9582 | 2048 | 1.1562 | 1.3724 | 0.2162 | 0.046741 | 1.0102 | 3 | 1.1562 | 1.0578 | 1.0255 | 0.5087
4 | 3433 | 512 | 1.6569 | 1.8835 | 0.22664 | 0.051367 | 1.6569 | 4 | 1.6569 | 1.5825 | 1.5342 | 0.61445
6 | 6113 | 512 | 2.9504 | 2.9059 | -0.044463 | 0.0019769 | 2.9504 | 5 | 2.3036 | 2.2163 | 2.1487 | 0.61445
8 | 8276 | 512 | 3.9943 | 3.9283 | -0.066044 | 0.0043619 | above uses | 6 | 2.9504 | 2.8501 | 2.7631 | 0.50719
12 | 12535 | 512 | 6.0499 | 5.973 | -0.076871 | 0.0059092 | ch 4, 6 to | 7 | 3.4723 | 3.3732 | 3.2703 | 0.50719
16 | 8435 | 256 | 8.1421 | 8.0177 | -0.12438 | 0.01547 | estimate | 8 | 3.9943 | 3.8963 | 3.7775 | 0.5015
24 | 12739 | 256 | 12.2966 | 12.1071 | -0.18947 | 0.035899 |  | 9 | 4.5082 | 4.4136 | 4.279 | 0.5015
32 | 4255 | 64 | 16.429 | 16.1966 | -0.23236 | 0.053991 |  | 10 | 5.0221 | 4.9309 | 4.7805 | 0.5015
48 | 6362 | 64 | 24.5643 | 24.3755 | -0.18879 | 0.035643 |  | 11 | 5.536 | 5.4482 | 5.282 | 0.5015
64 | 8442 | 64 | 32.5954 | 32.5544 | -0.040976 | 0.001679 |  | 12 | 6.0499 | 5.9655 | 5.7835 | 0.50786
128 | 8426 | 32 | 65.0672 | 65.2699 | 0.20279 | 0.041124 |  | 13 | 6.5729 | 6.4893 | 6.2914 | 0.50786
256 | 4213 | 8 | 130.1343 |  |  |  |  | 14 | 7.096 | 7.0131 | 6.7992 | 0.50786
512 | 8439 | 8 | 260.6702 |  |  |  |  | 15 | 7.619 | 7.537 | 7.3071 | 0.50786
1024 | 8455 | 4 | 522.3288 |  |  |  |  | 16 | 8.1421 | 8.0608 | 7.8149 | 0.515
2048 | 4239 | 1 | 1047.4993 |  |  |  |  | 17 | 8.6614 | 8.592 | 8.3299 | 0.515
2500 | 5182 | 1 | 1280.524 |  |  |  |  | 18 | 9.1807 | 9.1232 | 8.8449
3072 | 6383 | 1 | 1577.3031 |  |  |  |  | 19 | 9.7 | 6.9311 | 6.7197
3500 | 7284 | 1 | 1799.9492
4000 | 8343 | 1 | 2061.6387
4095 | 8546 | 1 | 2111.8021
HET 1 A1 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 21.3533 | =m | does not agree with tag1_f.c which has maybe erroneously used V2 HET 1 A1 AST HG 22C
0 | - | - | -0.054402 | 0.94142 | =b | is 1.862377 in tag1_f.c
8 | 6307 | 2048 | 0.32517 | 0.33056 | 0.0053892 | this pt scaled from HET 1 A1 AST HG 0C assuming same slope
12 | 3253 | 512 | 0.51496 | 0.51789 | 0.0029265 | this pt scaled from HET 1 A1 AST HG 0C assuming same slope
16 | 4797 | 512 | 0.69834 | 0.70521 | 0.0068739 | this pt scaled from HET 1 A1 AST HG 0C assuming same slope
24 | 7956 | 512 | 1.068 | 1.0799 | 0.011815
32 | 10828 | 512 | 1.4536 | 1.4545 | 0.00091624
48 | 8357 | 256 | 2.2438 | 2.2038 | -0.039945
64 | 10991 | 256 | 2.951 | 2.9531 | 0.0021568
128 | 5545 | 64 | 5.9551 | 5.9503 | -0.0047592
256 | 11102 | 64 | 11.923 | 11.9447 | 0.021682
512 | 11146 | 32 | 23.9406 | 23.9335 | -0.0070559 | is 23.950890 in tag1_f.c
1024 | 5588 | 8 | 48.01 |  |  | is 48.009979 in tag1_f.c
2048 | 11249 | 8 | 96.6472 |  |  | is 96.657156 fin tag1_f.c
3072 | 8455 | 4 | 145.2844 |  |  | is 145.284302 in tag1_f.c
4000 | 11048 | 4 | 189.8405 |  |  | is 189.840454 in tag1_f.c
4095 | 11315 | 4 | 194.4285 |  |  | is 194.438406 in tag1_f.c
HET 1 A1 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 21.6792 | =m | agrees with tag1_f.c
0 | - | - | -0.032185 | 0.31759 | =b
6 | 7555 | 2048 | 0.25355 | 0.26211 | 0.0085603
8 | 10393 | 2048 | 0.3488 | 0.35437 | 0.0055685
12 | 4012 | 512 | 0.53859 | 0.53888 | 0.00028981
16 | 5378 | 512 | 0.72196 | 0.72339 | 0.0014212
24 | 8132 | 512 | 1.0917 | 1.0924 | 0.00073072
32 | 10880 | 512 | 1.4606 | 1.4614 | 0.00084568
48 | 8244 | 256 | 2.2134 | 2.1995 | -0.01396
64 | 11010 | 256 | 2.9561 | 2.9375 | -0.018563
128 | 5482 | 64 | 5.8874 | 5.8896 | 0.0022251
256 | 10963 | 64 | 11.7737 | 11.7939 | 0.020174
512 | 10992 | 32 | 23.6098 | 23.6025 | -0.0072926
1024 | 5509 | 8 | 47.3313
2048 | 11072 | 8 | 95.1265
3072 | 8323 | 4 | 143.0162
4000 | 10872 | 4 | 186.8163
4095 | 11136 | 4 | 191.3527
HET 1 A1 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.4404 | =m | agrees with tag1_f.c
0 | - | - | -0.080815 | 0.58544 | =b
4 | 4486 | 512 | 0.8113 | 0.76898 | -0.042329 | this pt scaled from HET 1 A1 AST HG 22C assuming same slope as from ch 16 to 8
5 | 5729 | 512 | 1.0343 | 0.99418 | -0.040154 | this pt scaled from HET 1 A1 AST HG 22C assuming same slope as from ch 16 to 8
8 | 11463 | 512 | 1.7034 | 1.6698 | -0.033629 | this pt scaled from HET 1 A1 AST LG 0C assuming same slope
16 | 12585 | 256 | 3.4877 | 3.4714 | -0.01623 | this pt scaled from HET 1 A1 AST LG 0C assuming same slope
24 | 4908 | 64 | 5.271 | 5.2731 | 0.0021082
32 | 6606 | 64 | 7.0945 | 7.0747 | -0.019826
48 | 9935 | 64 | 10.6697 | 10.678 | 0.0082591
64 | 13264 | 64 | 14.2449 | 14.2813 | 0.036345
128 | 13335 | 32 | 28.6423 | 28.6944 | 0.052031
256 | 6679 | 8 | 57.3835 | 57.5206 | 0.1371
512 | 13415 | 8 | 115.2566 | 115.173 | -0.083676
1024 | 13462 | 4 | 231.3209
2048 | 6747 | 1 | 463.7416
3072 | 10166 | 1 | 698.7397
4000 | 13284 | 1 | 913.0492
4095 | 13606 | 1 | 935.1812
HET 1 A1 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.5133 | =m | agrees with tag1_f.c
0 | - | - | -0.12632 | 0.46083 | =b
8 | 12457 | 512 | 1.6723 | 1.6704 | -0.0018331
12 | 9578 | 256 | 2.5716 | 2.5567 | -0.014858
16 | 12874 | 256 | 3.4565 | 3.443 | -0.013519
24 | 4879 | 64 | 5.2398 | 5.2156 | -0.024265
32 | 6533 | 64 | 7.0161 | 6.9881 | -0.028031
48 | 9839 | 64 | 10.5666 | 10.5332 | -0.033414
64 | 13117 | 64 | 14.087 | 14.0783 | -0.0087271
128 | 13119 | 32 | 28.1784 | 28.2587 | 0.080357
256 | 6577 | 8 | 56.5071 | 56.6196 | 0.11247
512 | 13200 | 8 | 113.4095 | 113.3413 | -0.068175
1024 | 13241 | 4 | 227.5234
2048 | 6640 | 1 | 456.3871
3072 | 9997 | 1 | 687.1238
4000 | 13060 | 1 | 897.653
4095 | 13378 | 1 | 919.5101
HET 1 A2 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 21.4169 | =m | agrees with tag1_f.c
0 | - | - | -0.059906 | 1.2917 | =b
2 | 816 | 2048 | 0.027386 | 0.033072 | 0.0056867
4 | 3417 | 2048 | 0.11468 | 0.12646 | 0.011778
6 | 6252 | 2048 | 0.20982 | 0.21984 | 0.010017
8 | 9039 | 2048 | 0.30336 | 0.31322 | 0.0098665
12 | 3722 | 512 | 0.49966 | 0.49999 | 0.0003361
16 | 5114 | 512 | 0.68652 | 0.68676 | 0.00023637
24 | 7948 | 512 | 1.067 | 1.0603 | -0.0066753
32 | 10714 | 512 | 1.4383 | 1.4338 | -0.0044583
48 | 8178 | 256 | 2.1957 | 2.1809 | -0.014791
64 | 10957 | 256 | 2.9418 | 2.928 | -0.013848
128 | 5518 | 64 | 5.9261 | 5.9163 | -0.0098052
256 | 11063 | 64 | 11.8811 | 11.8928 | 0.011705
512 | 11102 | 32 | 23.8461 | 23.846 | -4.7751e-05
1024 | 5571 | 8 | 47.8639
2048 | 11213 | 8 | 96.3379
3072 | 8438 | 4 | 144.9923
4000 | 11033 | 4 | 189.5828
4095 | 11304 | 4 | 194.2395
HET 1 A2 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 21.5915 | =m | agrees with tag1_f.c
0 | - | - | -0.057591 | 0.44819 | =b
4 | 4494 | 2048 | 0.15082 | 0.1645 | 0.013678 | omitted pt at ch 3 due to high noise/signal ratio
6 | 7599 | 2048 | 0.25503 | 0.25713 | 0.0020997
8 | 10278 | 2048 | 0.34494 | 0.34976 | 0.004819
12 | 4011 | 512 | 0.53845 | 0.53502 | -0.0034355
16 | 5379 | 512 | 0.7221 | 0.72028 | -0.0018229
24 | 8157 | 512 | 1.095 | 1.0908 | -0.0042362
32 | 10924 | 512 | 1.4665 | 1.4613 | -0.0051727
48 | 8252 | 256 | 2.2156 | 2.2023 | -0.013221
64 | 10981 | 256 | 2.9483 | 2.9434 | -0.0048916
128 | 5497 | 64 | 5.9035 | 5.9075 | 0.0039938
256 | 11008 | 64 | 11.8221 | 11.8358 | 0.01371
512 | 11033 | 32 | 23.6978 | 23.6923 | -0.0055201
1024 | 5532 | 8 | 47.5289
2048 | 11080 | 8 | 95.1952
3072 | 8473 | 4 | 145.5937 |  |  | this pt not in tag1_f.c
4000 | 10954 | 4 | 188.2253
4095 | 11219 | 4 | 192.7789
HET 1 A2 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.4483 | =m | agrees with tag1_f.c
0 | - | - | -0.061887 | 1.0254 | =b
4 | 5147 | 512 | 0.69095 | 0.66871 | -0.022247 | omitted pt at ch 3 per note on cal sheet, as did tag1_f.c
6 | 7951 | 512 | 1.0674 | 1.1183 | 0.05094
8 | 11309 | 512 | 1.5182 | 1.5679 | 0.049756
16 | 12455 | 256 | 3.344 | 3.3664 | 0.022333
24 | 4882 | 64 | 5.243 | 5.1648 | -0.078253
32 | 6548 | 64 | 7.0322 | 6.9632 | -0.069028
48 | 9882 | 64 | 10.6128 | 10.5601 | -0.052725
64 | 13216 | 64 | 14.1934 | 14.1569 | -0.036422
128 | 13289 | 32 | 28.5435 | 28.5444 | 0.00086643
200 | 5193 | 8 | 44.6163 | 44.7303 | 0.11397
256 | 6660 | 8 | 57.2202 | 57.3193 | 0.099071
512 | 13379 | 8 | 114.9474 | 114.8691 | -0.078256
1024 | 13438 | 4 | 230.9085
2048 | 6747 | 1 | 463.7416
3072 | 10171 | 1 | 699.0833
4000 | 13300 | 1 | 914.1489
4095 | 13623 | 1 | 936.3497
HET 1 A2 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.5322 | =m | agrees with tag1_f.c
0 | - | - | -0.16965 | 0.73209 | =b
2 | 7435 | 2048 | 0.24953 | 0.27976 | 0.03023
3 | 3420 | 512 | 0.45911 | 0.5004 | 0.041286
4 | 5332 | 512 | 0.71579 | 0.72104 | 0.0052548
6 | 9054 | 512 | 1.2154 | 1.1623 | -0.053114
8 | 12078 | 512 | 1.6214 | 1.6036 | -0.01778
12 | 9357 | 256 | 2.5122 | 2.4862 | -0.026049
16 | 12637 | 256 | 3.3929 | 3.3688 | -0.024114
24 | 4809 | 64 | 5.1646 | 5.1339 | -0.030717
32 | 6446 | 64 | 6.9227 | 6.8991 | -0.023626
48 | 9714 | 64 | 10.4324 | 10.4294 | -0.0030016
64 | 12982 | 64 | 13.9421 | 13.9597 | 0.017623
128 | 13057 | 32 | 28.0452 | 28.0809 | 0.035684
256 | 6543 | 8 | 56.215 | 56.3233 | 0.10832
512 | 13137 | 8 | 112.8682 | 112.8082 | -0.059998
1024 | 13190 | 4 | 226.6471
2048 | 6551 | 1 | 450.2699 |  |  | this pt not in tag1_f.c
3072 | 9931 | 1 | 682.5874
4000 | 12999 | 1 | 893.4603
4095 | 13323 | 1 | 915.7298
HET 1 B1 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.2354 | =m | agrees with tag1_f.c
0 | - | - | -0.42243 | 2.366 | =b
3 | 1071 | 2048 | 0.12923 | 0.12111 | -0.008119
4 | 2595 | 2048 | 0.31311 | 0.31212 | -0.00099504
6 | 5845 | 2048 | 0.70525 | 0.69413 | -0.01112
8 | 9193 | 2048 | 1.1092 | 1.0761 | -0.03307
16 | 5480 | 512 | 2.6448 | 2.6042 | -0.040631
24 | 8557 | 512 | 4.1299 | 4.1323 | 0.0023635
32 | 11681 | 512 | 5.6377 | 5.6604 | 0.022675
48 | 9029 | 256 | 8.7155 | 8.7165 | 0.0010365
64 | 12186 | 256 | 11.7628 | 11.7726 | 0.0098045
128 | 6205 | 64 | 23.9581 | 23.9972 | 0.039085
256 | 12532 | 64 | 48.3872 | 48.4463 | 0.059035
512 | 12611 | 32 | 97.3845 | 97.3444 | -0.040064
1024 | 6325 | 8 | 195.3713
2048 | 12749 | 8 | 393.8007
3072 | 9601 | 4 | 593.1258
4000 | 12560 | 4 | 775.9254
4095 | 12866 | 4 | 794.8293
HET 1 B1 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.3465 | =m | agrees with tag1_f.c
0 | - | - | -0.24301 | 1.4913 | =b
2 | 640 | 2048 | 0.077222 | 0.095154 | 0.017932
3 | 1967 | 2048 | 0.23734 | 0.28219 | 0.044857
4 | 892 | 512 | 0.43051 | 0.46923 | 0.038721
6 | 1758 | 512 | 0.84848 | 0.84331 | -0.0051631
8 | 2587 | 512 | 1.2486 | 1.2174 | -0.03119
12 | 4114 | 512 | 1.9856 | 1.9655 | -0.020018
16 | 2848 | 256 | 2.7491 | 2.7137 | -0.035391
24 | 4409 | 256 | 4.2559 | 4.21 | -0.045867
32 | 1487 | 64 | 5.7414 | 5.7063 | -0.035106
48 | 2257 | 64 | 8.7145 | 8.699 | -0.015516
64 | 3026 | 64 | 11.6837 | 11.6916 | 0.007935
128 | 3055 | 32 | 23.5913 | 23.6621 | 0.070852
256 | 12319 | 64 | 47.5648 | 47.6032 | 0.03838
512 | 12369 | 32 | 95.5157 | 95.4853 | -0.030425
1000 | 6062 | 8 | 187.2476
1024 | 6208 | 8 | 191.7574
1536 | 9346 | 8 | 288.6863
2048 | 12496 | 8 | 385.9858
2500 | 15286 | 8 | 472.1654
3072 | 9402 | 4 | 580.8321
3500 | 10731 | 4 | 662.9344
4000 | 12295 | 4 | 759.5544
4090 | 12579 | 4 | 777.0992
4095 | 12592 | 4 | 777.9023
HET 1 B1 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.5638 | =m | agrees with tag1_f.c
0 | - | - | -1.516 | 2.265 | =b
3 | 3468 | 2048 | 0.41845 | 0.47005 | 0.051603
4 | 2203 | 512 | 1.0632 | 1.1095 | 0.046282
6 | 4994 | 512 | 2.4103 | 2.3885 | -0.021796
8 | 7848 | 512 | 3.7877 | 3.6675 | -0.12028
16 | 9200 | 256 | 8.8805 | 8.7833 | -0.097214
24 | 14400 | 256 | 13.8999 | 13.8991 | -0.00078819
32 | 4931 | 64 | 19.0391 | 19.015 | -0.024056
48 | 7581 | 64 | 29.271 | 29.2467 | -0.02426
64 | 10229 | 64 | 39.4951 | 39.4784 | -0.016741
128 | 10405 | 32 | 80.3494 | 80.4052 | 0.055807
256 | 5243 | 8 | 161.9497 | 162.2587 | 0.30901
512 | 10558 | 8 | 326.1234 | 325.9659 | -0.15757
1024 | 10598 | 4 | 654.7179
2048 | 5326 | 1 | 1316.1079
3072 | 8035 | 1 | 1985.5288
4000 | 10509 | 1 | 2596.879
4095 | 10767 | 1 | 2660.6334
HET 1 B1 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.5961 | =m | agrees with tag1_f.c
0 | - | - | -0.76981 | 1.4539 | =b
2 | 2144 | 2048 | 0.25869 | 0.34214 | 0.08345
3 | 6406 | 2048 | 0.77294 | 0.96868 | 0.19573
4 | 3048 | 512 | 1.4711 | 1.5952 | 0.12413
6 | 5986 | 512 | 2.8891 | 2.8483 | -0.040788
8 | 8695 | 512 | 4.1965 | 4.1013 | -0.095185
12 | 13801 | 512 | 6.6609 | 6.6075 | -0.053396
16 | 9568 | 256 | 9.2357 | 9.1136 | -0.12213
24 | 14763 | 256 | 14.2503 | 14.1259 | -0.12446
32 | 4986 | 64 | 19.2514 | 19.1381 | -0.11328
48 | 7574 | 64 | 29.2439 | 29.1627 | -0.081263
64 | 10147 | 64 | 39.1785 | 39.1872 | 0.0086702
128 | 10258 | 32 | 79.2142 | 79.2853 | 0.0711
256 | 5154 | 8 | 159.2006 | 159.4815 | 0.2809
512 | 10360 | 8 | 320.0075 | 319.874 | -0.13348
1024 | 10404 | 4 | 642.7331
2048 | 5223 | 1 | 1290.6555
3072 | 7871 | 1 | 1945.0028
4000 | 10291 | 1 | 2543.009
4095 | 10542 | 1 | 2605.0336
HET 1 B2 BST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.4725 | =m | does not agree with tag1_f.c | tag1_f.c uses fit to 14 pts but can't reproduce even if leave out ch 100 line
0 | - | - | -0.17194 | 1.0606 | =b | is -0.570460 in tag1_f.c
2 | 1169 | 2048 | 0.14105 | 0.17167 | 0.030616
3 | 2466 | 2048 | 0.29755 | 0.3544 | 0.056854
4 | 4023 | 2048 | 0.48541 | 0.53713 | 0.051721
6 | 7250 | 2048 | 0.87478 | 0.9026 | 0.02782
8 | 10046 | 2048 | 1.2121 | 1.2681 | 0.055923
10 | 12716 | 2048 | 1.5343 | 1.6335 | 0.099229
16 | 5537 | 512 | 2.6724 | 2.7299 | 0.05757 | omitted pts at ch 24, 32, 48; did line up with otehr cals
64 | 12275 | 256 | 11.8487 | 11.5011 | -0.34762
100 | 4634 | 64 | 17.8923 | 18.0795 | 0.1872
128 | 6071 | 64 | 23.4407 | 23.196 | -0.24467
256 | 12086 | 64 | 46.6652 | 46.5859 | -0.079316
512 | 12077 | 32 | 93.2609 | 93.3655 | 0.10467 | is 93.285118 in tag1_f.c
1000 | 5899 | 8 | 182.2127 |  |  | this pt not in tag1_f.c
1024 | 6048 | 8 | 186.8152
2048 | 12165 | 8 | 375.7616
3072 | 9154 | 4 | 565.5112
4000 | 11970 | 4 | 739.4767
4095 | 12262 | 4 | 757.5157
HET 1 B2 BST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.6177 | =m | agrees with tag1_f.c
0 | - | - | -0.012669 | 0.5319 | =b
2 | 2275 | 2048 | 0.2745 | 0.26134 | -0.013163
3 | 3465 | 2048 | 0.41808 | 0.43935 | 0.021262
4 | 5200 | 2048 | 0.62743 | 0.61736 | -0.010071
6 | 8407 | 2048 | 1.0144 | 0.97338 | -0.041005
8 | 10965 | 2048 | 1.323 | 1.3294 | 0.0063696
12 | 4294 | 512 | 2.0724 | 2.0414 | -0.031004
16 | 5747 | 512 | 2.7737 | 2.7535 | -0.020234
24 | 8666 | 512 | 4.1825 | 4.1776 | -0.0049689
32 | 11577 | 512 | 5.5875 | 5.6016 | 0.014157
48 | 8756 | 256 | 8.4519 | 8.4498 | -0.0021282
64 | 11681 | 256 | 11.2754 | 11.298 | 0.02261
100 | 4581 | 64 | 17.6877 | 17.7063 | 0.018667
128 | 5870 | 64 | 22.6646 | 22.6906 | 0.026003
256 | 11765 | 64 | 45.4258 | 45.4759 | 0.050162
512 | 11795 | 32 | 91.0832 | 91.0465 | -0.036657
1024 | 5917 | 8 | 182.7687
2048 | 11907 | 8 | 367.7923
3072 | 8960 | 4 | 553.5264
4000 | 11712 | 4 | 723.5381
4095 | 11996 | 4 | 741.0829
HET 1 B2 BST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.84691 | =m | agrees with tag1_f.c
0 | - | - | -0.71672 | 0.9697 | =b
2 | 2189 | 512 | 1.0565 | 1.2165 | 0.16005
3 | 4026 | 512 | 1.9431 | 2.3973 | 0.45421
4 | 6906 | 512 | 3.3331 | 3.5781 | 0.24498
6 | 12301 | 512 | 5.9369 | 5.9396 | 0.0026917
8 | 8372 | 256 | 8.0813 | 8.3011 | 0.21987
16 | 4501 | 64 | 17.3788 | 17.7473 | 0.36848
24 | 7180 | 64 | 27.7227 | 27.1934 | -0.52926
32 | 9613 | 64 | 37.1167 | 36.6395 | -0.47717
48 | 14488 | 64 | 55.9395 | 55.5318 | -0.40775
64 | 9681 | 32 | 74.7585 | 74.424 | -0.33446
128 | 4854 | 8 | 149.934 | 149.9931 | 0.059068
256 | 9740 | 8 | 300.8564 | 301.1311 | 0.27469
512 | 9768 | 4 | 603.4426 | 603.4072 | -0.035407
1024 | 4900 | 1 | 1210.839
2048 | 9866 | 1 | 2437.9873
3072 | 14874 | 1 | 3675.5141
3200 | 15503 | 1 | 3830.9463
4095 | - | - | 4917.7573
HET 1 B2 BST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.86522 | =m | agrees with tag1_f.c
0 | - | - | 0.17954 | 0.60183 | =b
6 | 13485 | 512 | 6.5084 | 6.239 | -0.26931
8 | 8928 | 256 | 8.618 | 8.5506 | -0.067369
12 | 13710 | 256 | 13.2339 | 13.1737 | -0.060219
16 | 4628 | 64 | 17.8691 | 17.7968 | -0.072374
24 | 7016 | 64 | 27.0894 | 27.0429 | -0.04649
32 | 9400 | 64 | 36.2943 | 36.2891 | -0.0051619
48 | 14169 | 64 | 54.7078 | 54.7815 | 0.073634
64 | 9476 | 32 | 73.1754 | 73.2738 | 0.098374
128 | 4757 | 8 | 146.9378 | 147.2432 | 0.30545
256 | 9548 | 8 | 294.9258 | 295.182 | 0.25626
512 | 9571 | 4 | 591.2725 | 591.0597 | -0.21279
1024 | 4799 | 1 | 1185.8809 | 1182.8149
2048 | 9662 | 1 | 2387.5768 | 2366.3254
3072 | 14559 | 1 | 3597.6745 | 3549.8359
4095 | - | - | 4806.5904 | 4732.1906
HET 1 C1 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0797 | =m | agrees with tag1_f.c ; fit thru ch 400 line
0 | - | - | -0.94657 | 1.5774 | =b
2 | 1785 | 2048 | 0.21538 | 0.39136 | 0.17599
3 | 1650 | 512 | 0.79635 | 1.3175 | 0.52117
4 | 3059 | 512 | 1.4764 | 2.2437 | 0.76729
6 | 7385 | 512 | 3.5643 | 4.096 | 0.53171
8 | 12124 | 512 | 5.8515 | 5.9483 | 0.096803
16 | 13908 | 256 | 13.425 | 13.3575 | -0.067495
32 | 7469 | 64 | 28.8385 | 28.176 | -0.66251
64 | 15191 | 64 | 58.6539 | 57.813 | -0.84092
100 | 11918 | 32 | 92.033 | 91.1545 | -0.87851
200 | 5952 | 8 | 183.8498 | 183.77 | -0.079857
400 | 11932 | 8 | 368.5646 | 369.0009 | 0.43635
600 | 8973 | 4 | 554.3295 | 554.2318 |  | tag1_f.c also has an interpolated pt for ch 512
800 | 12010 | 4 | 741.9478 | 739.4628
1000 | 15071 | 4 | 931.0487 | 924.6937
1200 | 4529 | 1 | 1119.1612 | 1109.9246
1400 | 5336 | 1 | 1318.579 | 1295.1556 |  | omitted pt at ch 1500; too high
4095 | - | - | 4005.7334 | 3791.1423
HET 1 C1 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1023 | =m | agrees with tag1_f.c
0 | - | - | -0.36886 | 1.1368 | =b
2 | 5701 | 2048 | 0.68788 | 0.78306 | 0.095185
3 | 2520 | 512 | 1.2162 | 1.6903 | 0.47403
4 | 4023 | 512 | 1.9416 | 2.5975 | 0.65584
6 | 7784 | 512 | 3.7568 | 4.4119 | 0.65506
8 | 12879 | 512 | 6.2159 | 6.2263 | 0.010444
12 | 10241 | 256 | 9.8854 | 9.8552 | -0.030203
16 | 14123 | 256 | 13.6326 | 13.484 | -0.14855
24 | 5453 | 64 | 21.0545 | 20.7417 | -0.31286
32 | 7364 | 64 | 28.4331 | 27.9994 | -0.43373
48 | 11158 | 64 | 43.0821 | 42.5147 | -0.56736
64 | 14906 | 64 | 57.5535 | 57.0301 | -0.52339
128 | 14934 | 32 | 115.3231 | 115.0915 | -0.23163
256 | 7476 | 8 | 230.9243 | 231.2144 | 0.29009
512 | 15002 | 8 | 463.393 | 463.4601 | 0.067097
1024 | 15107 | 4 | 933.2727
1200 | 4430 | 1 | 1094.6973
1400 | 5196 | 1 | 1283.9836
4095 | - | - | 3834.6159
HET 1 C1 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21802 | =m | agrees with tag1_f.c
0 | - | - | -5.81 | 1.5022 | =b
2 | 1153 | 256 | 1.113 | 2.2832 | 1.1702
3 | 4739 | 256 | 4.5744 | 6.8699 | 2.2954
4 | 2206 | 64 | 8.5176 | 11.4566 | 2.939
6 | 5067 | 64 | 19.5642 | 20.6301 | 1.0659
8 | 7556 | 64 | 29.1744 | 29.8035 | 0.62908
16 | 8654 | 32 | 66.8278 | 66.4973 | -0.33051
32 | 4626 | 8 | 142.8914 | 139.8849 | -3.0065
64 | 9414 | 8 | 290.7867 | 286.6601 | -4.1266
128 | 9460 | 4 | 584.4152 | 580.2105 | -4.2047
256 | 4709 | 1 | 1163.641 | 1167.3112 | 3.6702
512 | 9476 | 1 | 2341.6144 | 2341.5127 | -0.10162
800 | 14887 | 1 | 3678.7266
4095 | - | - | 18976.5902 |  |  | is 18977.76174 in tag1_f.c
HET 1 C1 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.22299 | =m | agrees with tag1_f.c
0 | - | - | -4.1594 | 1.0447 | =b
2 | 3025 | 256 | 2.92 | 4.284 | 1.3641
3 | 6692 | 256 | 6.4596 | 8.7686 | 2.309
4 | 2812 | 64 | 10.8574 | 13.2531 | 2.3957
6 | 5549 | 64 | 21.4252 | 22.2222 | 0.79703
8 | 8110 | 64 | 31.3135 | 31.1913 | -0.12213
12 | 12731 | 64 | 49.1556 | 49.1295 | -0.026042
16 | 8760 | 32 | 67.6464 | 67.0677 | -0.57862
24 | 13467 | 32 | 103.9947 | 102.9442 | -1.0505
32 | 4541 | 8 | 140.2658 | 138.8206 | -1.4453
48 | 6890 | 8 | 212.8235 | 210.5734 | -2.2501
64 | 9217 | 8 | 284.7016 | 282.3262 | -2.3754
128 | 9243 | 4 | 571.0094 | 569.3375 | -1.672
256 | 4614 | 1 | 1140.1655 | 1143.36 | 3.1945
512 | 9275 | 1 | 2291.9453 | 2291.4051 | -0.5402
800 | 14559 | 1 | 3597.6745 | 3582.9557
4095 | - | - | 18536.4865 | 18359.5514 |  | is 18542.67269 in tag1_f.c
HET 1 C1 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 3.8669 | =m | agrees with tag1_f.c
0 | - | - | -0.56167 | 1.6835 | =b
5 | 6660 | 2048 | 0.80359 | 0.85766 | 0.054074
6 | 8923 | 2048 | 1.0766 | 1.1163 | 0.039629
8 | 3337 | 512 | 1.6106 | 1.6335 | 0.022926
12 | 5475 | 512 | 2.6424 | 2.6679 | 0.025475
16 | 7635 | 512 | 3.6849 | 3.7023 | 0.017406
32 | 8173 | 256 | 7.8892 | 7.84 | -0.049137
64 | 4187 | 64 | 16.1664 | 16.1155 | -0.050946
128 | 8475 | 64 | 32.7228 | 32.6663 | -0.056494
256 | 8523 | 32 | 65.8162 | 65.7679 | -0.048286
512 | 4271 | 8 | 131.9259 | 131.9712 | 0.045354
1024 | 8549 | 8 | 264.0679
2048 | 8588 | 4 | 530.5452
3072 | 12946 | 4 | 799.7715
3500 | 14781 | 4 | 913.1332
4090 | 4320 | 1 | 1067.5152
4095 | 4326 | 1 | 1068.9979
HET 1 C1 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 3.9531 | =m | agrees wtih tag1_f.c
0 | - | - | -0.31456 | 0.81927 | =b
4 | 6169 | 2048 | 0.74435 | 0.80461 | 0.060261
6 | 10557 | 2048 | 1.2738 | 1.3105 | 0.036735
8 | 3746 | 512 | 1.808 | 1.8165 | 0.0085031
12 | 5842 | 512 | 2.8196 | 2.8283 | 0.0087488
16 | 7939 | 512 | 3.8317 | 3.8402 | 0.008512
24 | 12160 | 512 | 5.8689 | 5.8639 | -0.004993
32 | 8214 | 256 | 7.9288 | 7.8876 | -0.041182
48 | 12410 | 256 | 11.979 | 11.935 | -0.04406
64 | 4152 | 64 | 16.0313 | 15.9824 | -0.048868
128 | 8331 | 64 | 32.1668 | 32.172 | 0.005258
256 | 8359 | 32 | 64.5498 | 64.5513 | 0.0015391
512 | 4186 | 8 | 129.3003 | 129.3099 | 0.0095457
1024 | 8387 | 8 | 259.0639
2048 | 8415 | 4 | 519.8577
3072 | 12673 | 4 | 782.9063
3250 | 13417 | 4 | 828.8687
3800 | 15726 | 4 | 971.513
4095 | - | - | 1048.0222
HET 1 C1 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.77967 | =m | agrees wtih tag1_f.c; fit thru ch 256 line
0 | - | - | -3.0059 | 1.6102 | =b
5 | 8852 | 512 | 4.2723 | 4.3477 | 0.075424
6 | 11868 | 512 | 5.7279 | 5.6303 | -0.097614
8 | 8464 | 256 | 8.1701 | 8.1955 | 0.025433
12 | 13758 | 256 | 13.2802 | 13.3259 | 0.045653
16 | 4793 | 64 | 18.5062 | 18.4563 | -0.04996
256 | 10563 | 8 | 326.2779 | 326.2789 | 0.0010647
1024 | 5283 | 1 | 1305.4821 |  |  | tag1_f.c also has an interpolated pt for ch 512
2048 | 10640 | 1 | 2629.2504
3072 | 16038 | 1 | 3963.1502
4095 | - | - | 5295.7473
HET 1 C1 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.79707 | =m | agrees wtih tag1_f.c
0 | - | - | -1.4942 | 0.76297 | =b
4 | 7864 | 512 | 3.7955 | 4.0611 | 0.26569
6 | 13344 | 512 | 6.4403 | 6.5703 | 0.13002
8 | 9274 | 256 | 8.9519 | 9.0795 | 0.12756
12 | 14533 | 256 | 14.0283 | 14.0979 | 0.069543
16 | 4958 | 64 | 19.1433 | 19.1162 | -0.02708
24 | 7588 | 64 | 29.298 | 29.1529 | -0.14504
32 | 10214 | 64 | 39.4372 | 39.1897 | -0.24755
48 | 15420 | 64 | 59.5381 | 59.2631 | -0.27496
64 | 10294 | 32 | 79.4922 | 79.3365 | -0.15565
128 | 5162 | 8 | 159.4477 | 159.6303 | 0.18259
256 | 10363 | 8 | 320.1001 | 320.2179 | 0.11774
512 | 10383 | 4 | 641.4358 | 641.3929 | -0.042858
1024 | 5184 | 1 | 1281.0182
2048 | 10424 | 1 | 2575.8746
3072 | 15700 | 1 | 3879.627
4095 | - | - | 5182.1062
HET 1 C2 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0794 | =m | not in tag1_f.c but agrees with fit in book; C1 cal is used for C1+C2+C3
0 | - | - | -1.067 | 1.5236 | =b
2 | 1671 | 2048 | 0.20162 | 0.4414 | 0.23978
3 | 1732 | 512 | 0.83593 | 1.3679 | 0.53195
4 | 3186 | 512 | 1.5377 | 2.2943 | 0.75667
6 | 7709 | 512 | 3.7206 | 4.1473 | 0.42664
8 | 12352 | 512 | 5.9615 | 6.0002 | 0.038703
16 | 14017 | 256 | 13.5302 | 13.412 | -0.11824
24 | 5489 | 64 | 21.1935 | 20.8238 | -0.36977
32 | 7413 | 64 | 28.6223 | 28.2355 | -0.38674
48 | 11254 | 64 | 43.4527 | 43.0591 | -0.39366
64 | 15096 | 64 | 58.2871 | 57.8826 | -0.40445
128 | 15231 | 32 | 117.6166 | 117.1768 | -0.43985
256 | 7640 | 8 | 235.99 | 235.7651 | -0.22493
512 | 15300 | 8 | 472.5979 | 472.9418 | 0.3439
600 | 8951 | 4 | 552.9704
1024 | 15319 | 4 | 946.3695
2048 | 7728 | 1 | 1909.6661
2200 | 8316 | 1 | 2054.9668
2400 | 9092 | 1 | 2246.7241
2600 | 9886 | 1 | 2442.9295
4095 | - | - | 3909.5644
HET 1 C2 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1042 | =m | not in tag1_f.c but agrees with fit in book; C1 cal is used for C1+C2+C3
0 | - | - | -0.39576 | 1.1725 | =b
2 | 5752 | 2048 | 0.69403 | 0.74944 | 0.055405
3 | 2567 | 512 | 1.2389 | 1.6551 | 0.41618
4 | 4157 | 512 | 2.0063 | 2.5608 | 0.55445
6 | 8154 | 512 | 3.9354 | 4.3721 | 0.43669
8 | 13180 | 512 | 6.3612 | 6.1835 | -0.1777
12 | 10281 | 256 | 9.924 | 9.8061 | -0.11785
16 | 14112 | 256 | 13.6219 | 13.4288 | -0.19313
24 | 5415 | 64 | 20.9078 | 20.6742 | -0.23366
32 | 7298 | 64 | 28.1783 | 27.9195 | -0.25875
48 | 11032 | 64 | 42.5956 | 42.4102 | -0.18536
64 | 14769 | 64 | 57.0245 | 56.9009 | -0.12356
128 | 14906 | 32 | 115.1069 | 114.8638 | -0.24316
256 | 7475 | 8 | 230.8934 | 230.7894 | -0.10397
512 | 14972 | 8 | 462.4664 | 462.6408 | 0.17441
600 | 8759 | 4 | 541.1091 | 542.3397
1024 | 14988 | 4 | 925.9212 | 926.3435
2048 | 7560 | 1 | 1868.1516 | 1853.7488
2200 | 8126 | 1 | 2008.0159 | 1991.4105
2400 | 8859 | 1 | 2189.1475 | 2172.5444
2600 | 9621 | 1 | 2377.4453 | 2353.6783
2800 | 10437 | 1 | 2579.0871 | 2534.8121
4095 | - | - | 3884.7175 | 3707.6538
HET 1 C2 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21803 | =m | not in tag1_f.c but agrees with fit in book; C1 cal is used for C1+C2+C3
0 | - | - | -6.2173 | 1.7269 | =b
2 | 1163 | 256 | 1.1226 | 1.2524 | 0.12976
3 | 4965 | 256 | 4.7926 | 5.8389 | 1.0463
4 | 2330 | 64 | 8.9963 | 10.4255 | 1.4291
6 | 5281 | 64 | 20.3904 | 19.5986 | -0.79186
8 | 7668 | 64 | 29.6069 | 28.7717 | -0.83519
16 | 8750 | 32 | 67.5691 | 65.4641 | -2.1051
256 | 4711 | 1 | 1164.1352 | 1166.2362 | 2.101
512 | 9475 | 1 | 2341.3672 | 2340.3931 | -0.97411
800 | 14886 | 1 | 3678.4795 | 3661.3197
4095 | - | - | 18976.3431 | 18774.0037
HET 1 C2 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21976 | =m | not in tag1_f.c but agrees with fit in book; C1 cal is used for C1+C2+C3
0 | - | - | -4.3785 | 1.0522 | =b
2 | 3090 | 256 | 2.9827 | 4.313 | 1.3303
3 | 6903 | 256 | 6.6633 | 8.8634 | 2.2001
4 | 2973 | 64 | 11.479 | 13.4137 | 1.9347
6 | 5878 | 64 | 22.6955 | 22.5144 | -0.18116
8 | 8112 | 64 | 31.3212 | 31.615 | 0.29382
12 | 12914 | 64 | 49.8622 | 49.8163 | -0.045826
16 | 8838 | 32 | 68.2487 | 68.0177 | -0.23103
24 | 13589 | 32 | 104.9368 | 104.4203 | -0.51649
32 | 4581 | 8 | 141.5014 | 140.823 | -0.6784
48 | 6941 | 8 | 214.3988 | 213.6283 | -0.77055
64 | 9301 | 8 | 287.2963 | 286.4336 | -0.8627
128 | 9384 | 4 | 579.7201 | 577.6548 | -2.0653
256 | 4704 | 1 | 1162.4054 | 1160.0972 | -2.3083
512 | 9401 | 1 | 2323.0811 | 2324.982 | 1.9008
800 | 14712 | 1 | 3635.4823 | 3635.4774
4095 | - | - | 18650.6281 | 18628.8188
HET 1 C2 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1078 | =m | agrees with tag1_f.c
0 | - | - | -0.22298 | 0.40451 | =b
2 | 9560 | 2048 | 1.1535 | 1.4403 | 0.28676
3 | 3816 | 512 | 1.8417 | 2.343 | 0.50123
4 | 6044 | 512 | 2.9171 | 3.2457 | 0.32862
6 | 10450 | 512 | 5.0436 | 5.0511 | 0.0075429
8 | 13581 | 512 | 6.5547 | 6.8565 | 0.30182
16 | 14366 | 256 | 13.8671 | 14.0782 | 0.21106
24 | 5600 | 64 | 21.6221 | 21.2999 | -0.32227
32 | 7473 | 64 | 28.854 | 28.5215 | -0.33243
48 | 11220 | 64 | 43.3215 | 42.9649 | -0.35662
64 | 14968 | 64 | 57.7929 | 57.4082 | -0.38466
128 | 14966 | 32 | 115.5703 | 115.1815 | -0.38872
256 | 7473 | 8 | 230.8316 | 230.7282 | -0.10339
512 | 14943 | 8 | 461.5706 | 461.8216 | 0.25104
1024 | 14934 | 4 | 922.5852 | 924.0084
2048 | 7522 | 1 | 1858.7614 | 1848.382
2200 | 8093 | 1 | 1999.8612 | 1985.5937
2400 | 8847 | 1 | 2186.1822 | 2166.1354
2600 | 9609 | 1 | 2374.48 | 2346.6771 |  | omitted pt at ch 2800 not in tag1_f.c
4095 | - | - | 3782.0062 | 3696.2264 |  | this is 3782.005928 in tag1_f.c
HET 1 C2 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1284 | =m | agrees with tag1_f.c
0 | - | - | -0.24433 | -0.034088 | =b
2 | 12101 | 2048 | 1.4601 | 1.8026 | 0.34247
3 | 4791 | 512 | 2.3123 | 2.6887 | 0.37643
4 | 7177 | 512 | 3.4639 | 3.5749 | 0.11103
6 | 11392 | 512 | 5.4982 | 5.3473 | -0.15093
8 | 14421 | 512 | 6.9601 | 7.1196 | 0.15952
12 | 11124 | 256 | 10.7377 | 10.6643 | -0.073366
16 | 14790 | 256 | 14.2764 | 14.209 | -0.067349
24 | 5560 | 64 | 21.4677 | 21.2985 | -0.16922
32 | 7400 | 64 | 28.5721 | 28.3879 | -0.18421
48 | 11055 | 64 | 42.6844 | 42.5667 | -0.11767
64 | 14714 | 64 | 56.8121 | 56.7456 | -0.066579
128 | 14723 | 32 | 113.6938 | 113.4609 | -0.23287
256 | 7347 | 8 | 226.9396 | 226.8916 | -0.048054
512 | 14686 | 8 | 453.6322 | 453.753 | 0.12079
1024 | 14666 | 4 | 906.0288 | 907.4757
2048 | 7390 | 1 | 1826.1429 | 1814.9213 |  | omitted pts at ch 2900 and 2950 as did tag1_f.c
4095 | - | - | 3665.4725 | 3628.9262 |  | this is 3665.471325 in tag1_f.c
HET 1 C2 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.2142 | =m | agrees with tag1_f.c
0 | - | - | -4.2723 | 0.87011 | =b
2 | 5116 | 256 | 4.9383 | 5.275 | 0.33663
3 | 9887 | 256 | 9.5437 | 9.9436 | 0.39989
4 | 3720 | 64 | 14.3633 | 14.6121 | 0.24887
6 | 6524 | 64 | 25.1898 | 23.9493 | -1.2405
8 | 8651 | 64 | 33.4023 | 33.2865 | -0.11586
16 | 9099 | 32 | 70.2642 | 70.6351 | 0.37094 | fit is to ch 2-16; other data not taken; pt at ch 512 in tag1_f.c (extrapolated)
4095 | - | - | 19113.7766 | 19113.7828 |  | this is 19113.77657 in tag1_f.c
HET 1 C2 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.22455 | =m | agrees with tag1_f.c
0 | - | - | -0.76643 | -0.12182 | =b
2 | 7666 | 256 | 7.3998 | 9.4491 | 2.0493
3 | 11896 | 256 | 11.4829 | 13.9024 | 2.4195
4 | 4504 | 64 | 17.3904 | 18.3557 | 0.96529
6 | 7209 | 64 | 27.8346 | 27.2622 | -0.57239
8 | 9159 | 64 | 35.3638 | 36.1688 | 0.80505
12 | 14021 | 64 | 54.1364 | 53.982 | -0.15444
16 | 9323 | 32 | 71.994 | 71.7951 | -0.19886
24 | 14019 | 32 | 108.2573 | 107.4214 | -0.83596
32 | 4659 | 8 | 143.9107 | 143.0477 | -0.86301
48 | 6963 | 8 | 215.0784 | 214.3003 | -0.7781
64 | 9271 | 8 | 286.3696 | 285.5528 | -0.81675
128 | 9268 | 4 | 572.5539 | 570.5632 | -1.9907
256 | 4622 | 1 | 1142.1424 | 1140.5838 | -1.5586
512 | 9223 | 1 | 2279.0955 | 2280.6252 | 1.5297
800 | 14420 | 1 | 3563.3262 | 3563.1717
4095 | - | - | 18256.1736 | 18236.7506 |  | this is 18238.668320 in tag1_f.c
HET 1 C3 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0707 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -1.1043 | 1.6078 | =b
2 | 1088 | 2048 | 0.13128 | 0.36627 | 0.23499
3 | 1552 | 512 | 0.74905 | 1.3002 | 0.55116
4 | 2962 | 512 | 1.4296 | 2.2341 | 0.80458
6 | 7243 | 512 | 3.4957 | 4.102 | 0.60629
8 | 11944 | 512 | 5.7646 | 5.9699 | 0.20529
16 | 13813 | 256 | 13.3333 | 13.4414 | 0.10809
24 | 5530 | 64 | 21.3518 | 20.9129 | -0.43892
32 | 7491 | 64 | 28.9235 | 28.3844 | -0.53901
48 | 11400 | 64 | 44.0165 | 43.3275 | -0.68901
64 | 15280 | 64 | 58.9975 | 58.2705 | -0.72703
128 | 15352 | 32 | 118.551 | 118.0426 | -0.50844
256 | 7687 | 8 | 237.4418 | 237.5868 | 0.14494
512 | 15424 | 8 | 476.4281 | 476.6751 | 0.24706
600 | 9036 | 4 | 558.2215 | 558.8618
700 | 10553 | 4 | 651.938 | 652.2557
1024 | 15485 | 4 | 956.6246 | 954.8519
2048 | 7794 | 1 | 1925.9753 | 1911.2054
2500 | 9562 | 1 | 2362.8658 | 2333.3458 |  | omitted pt at ch 2700; too high
4095 | - | - | 3904.5479 | 3822.9784
HET 1 C3 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0966 | =m | not in tag1_f.c; does not agree with fit in book
0 | - | - | -0.47612 | 1.2327 | =b
2 | 5098 | 2048 | 0.61512 | 0.6997 | 0.084584
3 | 2405 | 512 | 1.1607 | 1.6116 | 0.45088
4 | 3916 | 512 | 1.89 | 2.5235 | 0.63353
6 | 7701 | 512 | 3.7168 | 4.3474 | 0.63059
8 | 12749 | 512 | 6.1531 | 6.1712 | 0.018073
12 | 10172 | 256 | 9.8188 | 9.8189 | 0.00011697
16 | 14073 | 256 | 13.5843 | 13.4665 | -0.11775
24 | 5445 | 64 | 21.0237 | 20.7619 | -0.26177
32 | 7364 | 64 | 28.4331 | 28.0572 | -0.37587
48 | 11194 | 64 | 43.2211 | 42.6479 | -0.57318
64 | 14992 | 64 | 57.8855 | 57.2386 | -0.64693
128 | 14996 | 32 | 115.8019 | 115.6013 | -0.20063
256 | 7512 | 8 | 232.0363 | 232.3267 | 0.29044
512 | 15077 | 8 | 465.7097 | 465.7776 | 0.067903
800 | 11798 | 4 | 728.8509 | 728.4098
1024 | 15135 | 4 | 935.0025 | 932.6793
1536 | 5677 | 1 | 1402.8435 | 1399.581
2048 | 7605 | 1 | 1879.2715 | 1866.4827
2200 | 8178 | 1 | 2020.8656 | 2005.0942
2400 | 8939 | 1 | 2208.9163 | 2187.4777
2600 | 9706 | 1 | 2398.4497 | 2369.8612
4095 | - | - | 3815.2116 | 3733.1777
HET 1 C3 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21307 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -6.2685 | 1.5151 | =b
2 | 1004 | 256 | 0.96913 | 2.2758 | 1.3067
3 | 4753 | 256 | 4.5879 | 6.9691 | 2.3811
4 | 2206 | 64 | 8.5176 | 11.6623 | 3.1447
6 | 5092 | 64 | 19.6607 | 21.0487 | 1.3881
8 | 7647 | 64 | 29.5258 | 30.4352 | 0.9094
16 | 8757 | 32 | 67.6232 | 67.981 | 0.35779
32 | 4714 | 8 | 145.6096 | 143.0726 | -3.4361
64 | 9611 | 8 | 296.8718 | 293.2557 | -2.537
100 | 15075 | 8 | 465.6479 | 462.2118 | -3.616
128 | 9650 | 4 | 596.1529 | 593.6221 | -2.5308
256 | 4827 | 1 | 1192.8 | 1194.3548 | 1.5548
512 | 9691 | 1 | 2394.743 | 2395.8202 | 1.0772
800 | 15192 | 1 | 3754.0951 | 3747.4688
4095 | - | - | 19306.4049 | 19211.643
HET 1 C3 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21892 | =m | not in tag1_f.c; does not agree with fit in book
0 | - | - | -4.2955 | 0.97039 | =b
2 | 2988 | 256 | 2.8842 | 4.7031 | 1.8189
3 | 6707 | 256 | 6.4741 | 9.271 | 2.7969
4 | 2826 | 64 | 10.9115 | 13.8389 | 2.9275
6 | 5610 | 64 | 21.6607 | 22.9747 | 1.314
8 | 8237 | 64 | 31.8038 | 32.1105 | 0.30664
12 | 12940 | 64 | 49.9626 | 50.382 | 0.41947
16 | 8917 | 32 | 68.8587 | 68.6536 | -0.20516
24 | 13734 | 32 | 106.0565 | 105.1967 | -0.85983
32 | 4633 | 8 | 143.1076 | 141.7398 | -1.3678
48 | 7038 | 8 | 217.395 | 214.826 | -2.569
64 | 9423 | 8 | 291.0647 | 287.9123 | -3.1524
128 | 9429 | 4 | 582.5 | 580.2572 | -2.2429
256 | 4718 | 1 | 1165.865 | 1164.947 | -0.91802
512 | 9473 | 1 | 2332.5948 | 2334.3266 | 1.7317
800 | 14845 | 1 | 3645.1659 | 3649.8786
4095 | - | - | 18662.2554 | 18701.0731
HET 1 C3 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1035 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -0.22406 | 0.3957 | =b
2 | 9749 | 2048 | 1.1763 | 1.4538 | 0.27748
3 | 3888 | 512 | 1.8765 | 2.36 | 0.48347
4 | 5797 | 512 | 2.7978 | 3.2661 | 0.4683
6 | 10072 | 512 | 4.8611 | 5.0785 | 0.21739
8 | 13904 | 512 | 6.7106 | 6.8909 | 0.18028
16 | 14540 | 256 | 14.0351 | 14.1403 | 0.10522
24 | 5635 | 64 | 21.7573 | 21.3897 | -0.36754
32 | 7534 | 64 | 29.0895 | 28.6392 | -0.45032
48 | 11309 | 64 | 43.6651 | 43.138 | -0.52709
64 | 15065 | 64 | 58.1674 | 57.6369 | -0.53049
128 | 15014 | 32 | 115.9409 | 115.6323 | -0.30857
256 | 7484 | 8 | 231.1714 | 231.6233 | 0.45186 | fit is thru ch 256
1024 | 15026 | 4 | 928.2687 | 927.5688
2048 | 7552 | 1 | 1866.1747 | 1855.4961
2200 | 8123 | 1 | 2007.2745 | 1993.2354
2400 | 8879 | 1 | 2194.0897 | 2174.4712
2600 | 9643 | 1 | 2382.8817 | 2355.707 |  | omitted pt at ch 2800 because it is not in line with expectations
4095 | - | - | 3794.1022 | 3710.4447
HET 1 C3 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1252 | =m | not in tag1_f.c; does not agree with fit in book but pretty close
0 | - | - | -0.067569 | -0.095019 | =b
2 | 12032 | 2048 | 1.4518 | 1.862 | 0.4102
3 | 4582 | 512 | 2.2114 | 2.7507 | 0.5393
4 | 6864 | 512 | 3.3128 | 3.6395 | 0.32668
6 | 11278 | 512 | 5.4432 | 5.417 | -0.026153
8 | 14708 | 512 | 7.0986 | 7.1945 | 0.095928
12 | 11155 | 256 | 10.7676 | 10.7496 | -0.018026
16 | 14879 | 256 | 14.3623 | 14.3046 | -0.057655
24 | 5594 | 64 | 21.599 | 21.4147 | -0.18421
32 | 7465 | 64 | 28.8231 | 28.5248 | -0.29822
48 | 11159 | 64 | 43.0859 | 42.745 | -0.3409
64 | 14906 | 64 | 57.5535 | 56.9652 | -0.58822
128 | 14769 | 32 | 114.049 | 113.846 | -0.20294
256 | 7358 | 8 | 227.2794 | 227.6076 | 0.32822
512 | 14734 | 8 | 455.1148 | 455.1308 | 0.01599
1024 | 14771 | 4 | 912.5155 | 910.1772
2048 | 7413 | 1 | 1831.8264 | 1820.27
2600 | 9456 | 1 | 2336.6722 | 2310.8669
2800 | 10214 | 1 | 2523.9815 | 2488.6194 |  | omitted pts at ch 2900 and 2950 due to preamp saturation
4095 | - | - | 3736.8098 | 3710.4447
HET 1 C3 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.20988 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -3.9962 | 1.0355 | =b
2 | 5024 | 256 | 4.8495 | 4.5953 | -0.25422
3 | 9606 | 256 | 9.2724 | 9.3598 | 0.08742
4 | 3354 | 64 | 12.9501 | 14.1244 | 1.1743
6 | 6309 | 64 | 24.3596 | 23.6534 | -0.70623
8 | 8760 | 64 | 33.8232 | 33.1825 | -0.64072
16 | 9189 | 32 | 70.9592 | 71.2987 | 0.33949 | fit is only thru ch 16; other data not taken
4095 | - | - | 19005.677 | 19505.7995
HET 1 C3 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.22384 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | 0.084944 | -0.12336 | =b
2 | 7616 | 256 | 7.3515 | 9.4862 | 2.1347
3 | 11380 | 256 | 10.9848 | 13.9538 | 2.969
4 | 4287 | 64 | 16.5525 | 18.4213 | 1.8688
6 | 7122 | 64 | 27.4987 | 27.3565 | -0.14226
8 | 9323 | 64 | 35.997 | 36.2916 | 0.29458
12 | 14064 | 64 | 54.3024 | 54.1618 | -0.14065
16 | 9380 | 32 | 72.4341 | 72.032 | -0.40212
24 | 14081 | 32 | 108.7361 | 107.7724 | -0.96369
32 | 4689 | 8 | 144.8373 | 143.5129 | -1.3245
48 | 7033 | 8 | 217.2406 | 214.9937 | -2.2468
64 | 9366 | 8 | 289.304 | 286.4746 | -2.8294
128 | 9286 | 4 | 573.6659 | 572.3981 | -1.2678
256 | 4622 | 1 | 1142.1424 | 1144.2451 | 2.1026
512 | 9259 | 1 | 2287.9915 | 2287.939 | -0.052496
800 | 14494 | 1 | 3581.6123 | 3574.5947
850 | 15406 | 1 | 3806.9767 | 3797.9724
4095 | - | - | 18433.121 | 18295.1865
HET 1 C4 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1131 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -0.32868 | 0.52888 | =b
2 | 8836 | 2048 | 1.0661 | 1.3217 | 0.25553
3 | 3654 | 512 | 1.7636 | 2.2201 | 0.45653
4 | 5830 | 512 | 2.8138 | 3.1185 | 0.30472
6 | 10189 | 512 | 4.9176 | 4.9153 | -0.0022655
8 | 13419 | 512 | 6.4765 | 6.7121 | 0.23564
16 | 14225 | 256 | 13.731 | 13.8994 | 0.16843
24 | 5573 | 64 | 21.5179 | 21.0867 | -0.43114
32 | 7431 | 64 | 28.6918 | 28.274 | -0.41776
48 | 11144 | 64 | 43.028 | 42.6486 | -0.3794
64 | 14859 | 64 | 57.372 | 57.0232 | -0.34877
128 | 14853 | 32 | 114.6977 | 114.5216 | -0.17606
256 | 7419 | 8 | 229.1636 | 229.5183 | 0.35469
512 | 14877 | 8 | 459.5319 | 459.5118 | -0.020133
1024 | 14927 | 4 | 922.1527 | 919.4988
2048 | 7509 | 1 | 1855.549 | 1839.4727
2200 | 8078 | 1 | 1996.1546 | 1976.0313
2300 | 8454 | 1 | 2089.0679 | 2065.8725
2400 | 8832 | 1 | 2182.4755 | 2155.7137
2500 | 9212 | 1 | 2276.3773 | 2245.5549
2600 | 9598 | 1 | 2371.7618 | 2335.3961
2800 | 14363 | 1 | 3549.2409 | 2515.0785
4095 | - | - | 11173.4184 | 3678.5221
HET 1 C4 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1345 | =m | not in tag1_f.c; does not agree with fit in book but pretty close
0 | - | - | -0.1723 | 0.052289 | =b
2 | 11988 | 2048 | 1.4465 | 1.7168 | 0.27034
3 | 4674 | 512 | 2.2558 | 2.5982 | 0.3424
4 | 7005 | 512 | 3.3809 | 3.4797 | 0.098824
6 | 11311 | 512 | 5.4591 | 5.2426 | -0.21652
8 | 14372 | 512 | 6.9365 | 7.0055 | 0.069024
12 | 11022 | 256 | 10.6392 | 10.5313 | -0.10798
16 | 14684 | 256 | 14.1741 | 14.057 | -0.11703
24 | 5519 | 64 | 21.3094 | 21.1086 | -0.20076
32 | 7355 | 64 | 28.3983 | 28.1602 | -0.23816
48 | 10988 | 64 | 42.4257 | 42.2633 | -0.16237
64 | 14606 | 64 | 56.3951 | 56.3665 | -0.028673
128 | 14594 | 32 | 112.6976 | 112.779 | 0.08141
256 | 7292 | 8 | 225.2408 | 225.6041 | 0.36335
512 | 14614 | 8 | 451.4082 | 451.2543 | -0.15387
800 | 11430 | 4 | 706.1168 | 705.1108
1024 | 14660 | 4 | 905.6581 | 902.5547
2048 | 7364 | 1 | 1819.718 | 1805.1556
2500 | 9026 | 1 | 2230.4149 | 2203.5692 |  | omited pts at ch 2700 and 2900; preamp saturation
4095 | - | - | 3679.6658 | 3609.4758
HET 1 C4 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21412 | =m | not in tag1_f.c but does agree with fit in book
0 | - | - | -4.225 | 0.97286 | =b
2 | 4965 | 256 | 4.7926 | 4.7971 | 0.0044905
3 | 9636 | 256 | 9.3014 | 9.4674 | 0.16602
4 | 3416 | 64 | 13.1895 | 14.1377 | 0.94821
6 | 6382 | 64 | 24.6415 | 23.4783 | -1.1632
8 | 8581 | 64 | 33.132 | 32.819 | -0.31307
16 | 9042 | 32 | 69.824 | 70.1815 | 0.3575
4095 | - | - | 19017.1976 | 19120.4077
HET 1 C4 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.22619 | =m | not in tag1_f.c but does agree with fit in book
0 | - | - | -1.7336 | -0.022504 | =b
8 | 9079 | 64 | 35.0549 | 35.4686 | 0.41371
12 | 13843 | 64 | 53.4491 | 53.1531 | -0.296
16 | 9216 | 32 | 71.1677 | 70.8377 | -0.33001
24 | 13864 | 32 | 107.0604 | 106.2068 | -0.85365
32 | 4611 | 8 | 142.428 | 141.5758 | -0.85218
48 | 6856 | 8 | 211.7733 | 212.314 | 0.54075
64 | 9170 | 8 | 283.2498 | 283.0522 | -0.19764
128 | 9150 | 4 | 565.2641 | 566.0049 | 0.74077
256 | 4573 | 1 | 1130.034 | 1131.9103 | 1.8763
512 | 9165 | 1 | 2264.7632 | 2263.7211 | -1.042
800 | 14357 | 1 | 3547.7583 | 3537.0083
4095 | - | - | 18164.8734 | 18104.6514
HET 2 A1 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 20.3115 | =m | agrees with tag1_f.c
0 | - | - | -0.10307 | 1.3715 | =b
6 | 6229 | 2048 | 0.20905 | 0.22788 | 0.018825
8 | 9329 | 2048 | 0.31309 | 0.32634 | 0.013252
12 | 3846 | 512 | 0.5163 | 0.52328 | 0.0069732
16 | 5352 | 512 | 0.71847 | 0.72021 | 0.0017347
32 | 11416 | 512 | 1.5325 | 1.5079 | -0.024589
48 | 8615 | 256 | 2.313 | 2.2957 | -0.017352
64 | 11527 | 256 | 3.0949 | 3.0834 | -0.011458
128 | 5803 | 64 | 6.2321 | 6.2343 | 0.0021866
256 | 11660 | 64 | 12.5223 | 12.5362 | 0.013903
512 | 11706 | 32 | 25.1434 | 25.1399 | -0.003475
1024 | 5877 | 8 | 50.493 | 50.3474
2048 | 11835 | 8 | 101.6819 | 100.7622
3072 | 8908 | 4 | 153.0684 | 151.1771
3500 | 10168 | 4 | 174.7193 | 172.249
4000 | 11653 | 4 | 200.2364 | 196.8656
4090 | 11923 | 4 | 204.8759 | 201.2966 |  | note says to just use ch 4090 in fit (?)
4095 | 11923 | 4 | 205.1335 | 201.5428 |  | forced to be 205.133541 from tag1_f.c
HET 2 A1 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 20.6233 | =m | agrees with tag1_f.c
0 | - | - | -0.069606 | 0.86699 | =b
6 | 7079 | 2048 | 0.23758 | 0.24889 | 0.011316
8 | 10130 | 2048 | 0.33997 | 0.34587 | 0.0058989
12 | 4036 | 512 | 0.54181 | 0.53983 | -0.0019813
16 | 5478 | 512 | 0.73539 | 0.73378 | -0.0016056
24 | 8378 | 512 | 1.1247 | 1.1217 | -0.0030019
32 | 11300 | 512 | 1.517 | 1.5096 | -0.0073517
48 | 8551 | 256 | 2.2958 | 2.2854 | -0.010413
64 | 11415 | 256 | 3.0648 | 3.0613 | -0.00354
128 | 5736 | 64 | 6.1602 | 6.1645 | 0.0043518
256 | 11509 | 64 | 12.3601 | 12.3711 | 0.011007
512 | 11541 | 32 | 24.789 | 24.7843 | -0.0046799
1024 | 5790 | 8 | 49.7455 | 49.6107
2048 | 11654 | 8 | 100.1268 | 99.2633
3072 | 8770 | 4 | 150.6971 | 148.916
4000 | 11467 | 4 | 197.0403 | 193.9138
4095 | 11747 | 4 | 201.8516 | 198.5202
HET 2 A1 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.1301 | =m | does not agree with tag1_f.c | tag1_f.c uses fit to 11 pts but can't reproduce
0 | - | - | -0.27654 | 1.2931 | =b | is 1.260612 in tag1_f.c
6 | 8272 | 512 | 1.1105 | 1.1397 | 0.029204
8 | 11716 | 512 | 1.5728 | 1.6239 | 0.051122
12 | 9593 | 256 | 2.5756 | 2.5924 | 0.016827
16 | 13111 | 256 | 3.5201 | 3.5609 | 0.040794
24 | 5195 | 64 | 5.5792 | 5.498 | -0.062287
32 | 6981 | 64 | 7.4973 | 7.435 | -0.038373
48 | 10566 | 64 | 11.3474 | 11.309 | -0.016606
64 | 14153 | 64 | 15.1997 | 15.1831 | 0.0027992
128 | 14282 | 32 | 30.6764 | 30.6792 | 0.095308
256 | 7167 | 8 | 61.5762 | 61.6715 | -0.037564
512 | 14397 | 8 | 123.6936 | 123.6561 | -0.84458 | is 123.652306 in tag1_f.c
1024 | 14460 | 4 | 248.4698 | 247.6252
2048 | 7263 | 1 | 499.2078 | 495.5635
3072 | 10949 | 1 | 752.5576 | 743.5018
3500 | 12499 | 1 | 859.0938 | 847.1323
4000 | 14323 | 1 | 984.4628 | 968.1959
4090 | 14656 | 1 | 1007.3508 | 989.9874
4095 | 14673 | 1 | 1008.5193 | 991.198
HET 2 A1 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.2009 | =m | agrees with tag1_f.c
0 | - | - | -0.060276 | 0.84717 | =b
6 | 9145 | 512 | 1.2277 | 1.2266 | -0.0010641
8 | 12343 | 512 | 1.657 | 1.7027 | 0.04571
12 | 9913 | 256 | 2.6615 | 2.6549 | -0.0066645
16 | 13372 | 256 | 3.5902 | 3.607 | 0.016809
24 | 5131 | 64 | 5.5105 | 5.5114 | -0.037504
32 | 6940 | 64 | 7.4532 | 7.4157 | -0.046711
48 | 10495 | 64 | 11.2711 | 11.2244 | -0.021551
64 | 14018 | 64 | 15.0547 | 15.0331 | 0.0017637
128 | 14091 | 32 | 30.2661 | 30.2679 | 0.080612
256 | 7060 | 8 | 60.6569 | 60.7375 | -0.032329
512 | 14166 | 8 | 121.709 | 121.6766 | -0.73934
1024 | 14217 | 4 | 244.2943 | 243.5549
2048 | 7136 | 1 | 490.4787 | 487.3115
3072 | 10753 | 1 | 739.0859 | 731.0681
4000 | 14059 | 1 | 966.3172 | 951.9725
4095 | 14401 | 1 | 989.8239 | 974.5866
HET 2 A2 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 20.2228 | =m | agrees with tag1_f.c
0 | - | - | -0.064269 | 1.1451 | =b
2 | 766 | 2048 | 0.025708 | 0.042273 | 0.016565
4 | 3447 | 2048 | 0.11568 | 0.14117 | 0.025486
6 | 6627 | 2048 | 0.22241 | 0.24007 | 0.01766
8 | 9669 | 2048 | 0.3245 | 0.33897 | 0.014465
12 | 3939 | 512 | 0.52879 | 0.53676 | 0.002795
16 | 5451 | 512 | 0.73176 | 0.73456 | -0.02187
24 | 8509 | 512 | 1.1423 | 1.1302 | -0.026216
32 | 11429 | 512 | 1.5343 | 1.5257 | -1.3986
48 | 8711 | 256 | 2.3388 | 2.3169 | -0.026216
64 | 11674 | 256 | 3.1343 | 3.1081 | -1.3986
100 | 4570 | 64 | 4.908 | 4.8883 | -6.3203
128 | 5854 | 64 | 6.2869 | 6.2729 | -12.6506
256 | 11726 | 64 | 12.5932 | 12.6023 | -25.3777
512 | 11757 | 32 | 25.2529 | 25.2613 | -51.2487
1024 | 5894 | 8 | 50.639 | 50.5792
2048 | 11852 | 8 | 101.8279 | 101.2151
3072 | 8914 | 4 | 153.1715 | 151.851
3500 | 10171 | 4 | 174.7708 | 173.0152
4000 | 11657 | 4 | 200.3051 | 197.7398
4095 | 11941 | 4 | 205.1852 | 202.4374
HET 2 A2 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 20.5144 | =m | agrees with tag1_f.c
0 | - | - | -0.029836 | 0.66987 | =b
2 | 1603 | 2048 | 0.053798 | 0.064839 | 0.01104
3 | 2849 | 2048 | 0.095615 | 0.11358 | 0.01797
4 | 4369 | 2048 | 0.14663 | 0.16233 | 0.015703
6 | 7536 | 2048 | 0.25292 | 0.25982 | 0.0069074
8 | 10337 | 2048 | 0.34692 | 0.35732 | 0.010395
12 | 4137 | 512 | 0.55537 | 0.5523 | -0.0030679
16 | 5601 | 512 | 0.7519 | 0.74728 | -0.0046167
24 | 8481 | 512 | 1.1385 | 1.1373 | -0.0012708
32 | 11404 | 512 | 1.5309 | 1.5272 | -0.0036973
48 | 8662 | 256 | 2.3256 | 2.3072 | -0.018484
64 | 11588 | 256 | 3.1112 | 3.0871 | -0.024143
128 | 5796 | 64 | 6.2246 | 6.2069 | -0.01778
256 | 11586 | 64 | 12.4428 | 12.4464 | 0.0035366
512 | 11601 | 32 | 24.9179 | 24.9254 | 0.0075081
1024 | 5810 | 8 | 49.9173 | 49.8834
2048 | 11677 | 8 | 100.3244 | 99.7994
3072 | 8777 | 4 | 150.8174 | 149.7155
4000 | 11474 | 4 | 197.1606 | 194.9519
4095 | 11753 | 4 | 201.9547 | 199.5828
HET 2 A2 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.2713 | =m | agrees with tag1_f.c
0 | - | - | -0.31809 | 1.0157 | =b
2 | 4063 | 2048 | 0.13636 | 0.23043 | 0.094075
4 | 4401 | 512 | 0.59081 | 0.69867 | 0.10786
6 | 8118 | 512 | 1.0898 | 1.1669 | 0.077118
8 | 11562 | 512 | 1.5521 | 1.6352 | 0.08302
12 | 9394 | 256 | 2.5222 | 2.5716 | 0.04945
16 | 12902 | 256 | 3.464 | 3.5081 | 0.044071
24 | 5114 | 64 | 5.4922 | 5.3811 | -0.11113
32 | 6875 | 64 | 7.3834 | 7.254 | -0.12941
48 | 10348 | 64 | 11.1133 | 10.9999 | -0.11333
64 | 13839 | 64 | 14.8624 | 14.7458 | -0.11659
128 | 13872 | 32 | 29.7958 | 29.7295 | -0.066259
256 | 6941 | 8 | 59.6345 | 59.6968 | 0.062329
512 | 13922 | 8 | 119.6126 | 119.6314 | 0.018799
1024 | 13965 | 4 | 239.9641 | 239.5006
2048 | 7005 | 1 | 481.4747 | 479.239
3072 | 10549 | 1 | 725.0644 | 718.9774
4000 | 13797 | 1 | 948.3092 | 936.2404
4095 | 14136 | 1 | 971.6097 | 958.4817
HET 2 A2 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 4.3503 | =m | agrees with tag1_f.c
0 | - | - | -0.26809 | 0.71795 | =b
2 | 6444 | 2048 | 0.21627 | 0.2947 | 0.078437
3 | 3415 | 512 | 0.45844 | 0.52457 | 0.06613
4 | 5391 | 512 | 0.72371 | 0.75444 | 0.030733
6 | 8729 | 512 | 1.1718 | 1.2142 | 0.042365
8 | 12310 | 512 | 1.6525 | 1.6739 | 0.021375
12 | 9750 | 256 | 2.6178 | 2.5934 | -0.024362
16 | 13161 | 256 | 3.5336 | 3.5129 | -0.020698
24 | 4997 | 64 | 5.3665 | 5.3518 | -0.014712
32 | 6715 | 64 | 7.2116 | 7.1908 | -0.020808
48 | 10182 | 64 | 10.935 | 10.8687 | -0.066292
64 | 13628 | 64 | 14.6358 | 14.5466 | -0.089224
128 | 13649 | 32 | 29.3168 | 29.2583 | -0.058519
256 | 6825 | 8 | 58.6378 | 58.6815 | 0.043702
512 | 13678 | 8 | 117.5162 | 117.5281 | 0.011872
1024 | 13711 | 4 | 235.5995 | 235.2213
2048 | 6875 | 1 | 472.5394 | 470.6076
3072 | 10349 | 1 | 711.3178 | 705.9939
4000 | 13529 | 1 | 929.8888 | 919.3127
4095 | 13859 | 1 | 952.5706 | 941.1503
HET 2 B1 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.3365 | =m | agrees with tag1_f.c
0 | - | - | -0.38309 | 0.76092 | =b
4 | 3893 | 2048 | 0.46973 | 0.60696 | 0.13724
6 | 7427 | 2048 | 0.89614 | 0.98174 | 0.085605
8 | 10713 | 2048 | 1.2926 | 1.3565 | 0.063895
12 | 4260 | 512 | 2.056 | 2.1061 | 0.050036
16 | 5847 | 512 | 2.822 | 2.8556 | 0.033644
24 | 9183 | 512 | 4.4321 | 4.3547 | -0.077328
32 | 12296 | 512 | 5.9345 | 5.8538 | -0.080672
48 | 9297 | 256 | 8.9741 | 8.852 | -0.12211
64 | 12399 | 256 | 11.9684 | 11.8502 | -0.11818
128 | 6185 | 64 | 23.8809 | 23.8431 | -0.037783
256 | 12374 | 64 | 47.7772 | 47.8288 | 0.051577
512 | 12404 | 32 | 95.786 | 95.8001 | 0.014075
1024 | 6221 | 8 | 192.1589 | 191.7428
2048 | 12526 | 8 | 386.9125 | 383.6281
3072 | 9426 | 4 | 582.3147 | 575.5135
4000 | 12327 | 4 | 761.5312 | 749.4096
4095 | 12627 | 4 | 780.0645 | 767.2114
HET 2 B1 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.4352 | =m | agrees with tag1_f.c
0 | - | - | -0.056107 | 0.50276 | =b
2 | 1989 | 2048 | 0.23999 | 0.27547 | 0.035482
3 | 3216 | 2048 | 0.38804 | 0.45946 | 0.07142
4 | 4813 | 2048 | 0.58073 | 0.64345 | 0.062714
6 | 8247 | 2048 | 0.99508 | 1.0114 | 0.016345
8 | 11318 | 2048 | 1.3656 | 1.3794 | 0.013775
12 | 4444 | 512 | 2.1448 | 2.1153 | -0.029493
16 | 5997 | 512 | 2.8944 | 2.8513 | -0.043079
24 | 8985 | 512 | 4.3365 | 4.3232 | -0.013301
32 | 12033 | 512 | 5.8076 | 5.7951 | -0.01248
48 | 9108 | 256 | 8.7917 | 8.7389 | -0.052828
64 | 12177 | 256 | 11.7541 | 11.6827 | -0.071458
128 | 6081 | 64 | 23.4793 | 23.4579 | -0.021455
256 | 12163 | 64 | 46.9625 | 47.0082 | 0.04573
512 | 12187 | 32 | 94.1103 | 94.1089 | -0.0013715
1024 | 6113 | 8 | 188.8229 | 188.3104
2048 | 12300 | 8 | 379.9316 | 376.7132
3072 | 9249 | 4 | 571.3801 | 565.1161
4000 | 12091 | 4 | 746.9518 | 735.8562
4095 | 12387 | 4 | 765.2379 | 753.3349
HET 2 B1 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.6778 | =m | agrees with tag1_f.c
0 | - | - | -1.2756 | 0.79707 | =b
4 | 3087 | 512 | 1.4899 | 1.9091 | 0.41916
6 | 5952 | 512 | 2.8727 | 3.1011 | 0.22847
8 | 8477 | 512 | 4.0913 | 4.2932 | 0.20188
12 | 13452 | 512 | 6.4924 | 6.6773 | 0.1849
16 | 9271 | 256 | 8.9491 | 9.0615 | 0.11242
24 | 14595 | 256 | 14.0882 | 13.8298 | -0.25842
32 | 4906 | 64 | 18.9425 | 18.598 | -0.3445
48 | 7382 | 64 | 28.5026 | 28.1346 | -0.36801
64 | 9844 | 64 | 38.0086 | 37.6711 | -0.33747
128 | 9832 | 32 | 75.9245 | 75.8174 | -0.10719
256 | 4915 | 8 | 151.8182 | 152.1098 | 0.29159
512 | 9865 | 8 | 304.7175 | 304.6947 | -0.022845
1024 | 9905 | 4 | 611.9061 | 609.8644
2048 | 4975 | 1 | 1229.3722 | 1220.2039
3072 | 7496 | 1 | 1852.3366 | 1830.5435
4000 | 9804 | 1 | 2422.6664 | 2383.6636
4095 | 10044 | 1 | 2481.9728 | 2440.2869
HET 2 B1 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.7068 | =m | agrees with tag1_f.c
0 | - | - | -0.19402 | 0.42489 | =b
2 | 6214 | 2048 | 0.74978 | 0.92281 | 0.17304
3 | 10125 | 2048 | 1.2217 | 1.5087 | 0.28702
4 | 3902 | 512 | 1.8832 | 2.0946 | 0.21132
6 | 6727 | 512 | 3.2467 | 3.2663 | 0.019617
8 | 9196 | 512 | 4.4383 | 4.4381 | -0.00026269
12 | 14129 | 512 | 6.8192 | 6.7816 | -0.03761
16 | 9575 | 256 | 9.2425 | 9.1251 | -0.11743
24 | 14383 | 256 | 13.8835 | 13.8121 | -0.071464
32 | 4815 | 64 | 18.5912 | 18.4991 | -0.092103
48 | 7271 | 64 | 28.074 | 27.8731 | -0.20095
64 | 9715 | 64 | 37.5105 | 37.2471 | -0.26346
128 | 9694 | 32 | 74.8589 | 74.7431 | -0.11583
256 | 4840 | 8 | 149.5016 | 149.7351 | 0.2335
512 | 9704 | 8 | 299.7444 | 299.719 | -0.025386
1024 | 9737 | 4 | 601.5275 | 599.687
2048 | 4886 | 1 | 1207.3795 | 1199.623
3072 | 7363 | 1 | 1819.4709 | 1799.5589
4000 | 9623 | 1 | 2377.9395 | 2343.2509
4095 | 9857 | 1 | 2435.7633 | 2398.909
HET 2 B2 BST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.1786 | =m | agrees with tag1_f.c
0 | - | - | -0.07324 | 0.14098 | =b
2 | 2198 | 2048 | 0.26521 | 0.35898 | 0.093775
4 | 5003 | 2048 | 0.60366 | 0.74519 | 0.14153
6 | 8429 | 2048 | 1.017 | 1.1314 | 0.11436
8 | 11976 | 2048 | 1.445 | 1.5176 | 0.072594
12 | 4674 | 512 | 2.2558 | 2.29 | 0.034181
16 | 6291 | 512 | 3.0363 | 3.0624 | 0.026174
24 | 9647 | 512 | 4.656 | 4.6073 | -0.048723
32 | 12861 | 512 | 6.2072 | 6.1521 | -0.055084
48 | 9665 | 256 | 9.3294 | 9.2418 | -0.087596
64 | 12902 | 256 | 12.454 | 12.3314 | -0.12252
128 | 6453 | 64 | 24.9156 | 24.6901 | -0.22554
256 | 12809 | 64 | 49.4567 | 49.4074 | -0.049331
512 | 12786 | 32 | 98.7359 | 98.8421 | 0.10617
1024 | 6441 | 8 | 198.9544 | 197.7113
2048 | 13011 | 8 | 401.8935 | 395.4499
3072 | 9795 | 4 | 605.1106 | 593.1885
4000 | 12807 | 4 | 791.1844 | 772.3891
4095 | 13119 | 4 | 810.459 | 790.734
HET 2 B2 BST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 5.2198 | =m | agrees with tag1_f.c
0 | - | - | -0.091701 | 0.39279 | =b
2 | 2132 | 2048 | 0.25725 | 0.30791 | 0.050663
3 | 3578 | 2048 | 0.43172 | 0.49949 | 0.067768
4 | 5318 | 2048 | 0.64167 | 0.69107 | 0.0494
6 | 8712 | 2048 | 1.0512 | 1.0742 | 0.023041
8 | 11679 | 2048 | 1.4092 | 1.4574 | 0.048203
12 | 4632 | 512 | 2.2356 | 2.2237 | -0.011877
16 | 6202 | 512 | 2.9933 | 2.99 | -0.0033007
24 | 9373 | 512 | 4.5238 | 4.5226 | -0.0011107
32 | 12582 | 512 | 6.0725 | 6.0553 | -0.017261
48 | 9510 | 256 | 9.1798 | 9.1205 | -0.059214
64 | 12676 | 256 | 12.2358 | 12.1858 | -0.050008
128 | 6352 | 64 | 24.5257 | 24.4468 | -0.078821
256 | 12708 | 64 | 49.0668 | 48.9689 | -0.097837
512 | 12682 | 32 | 97.9328 | 98.0131 | 0.080354
2048 | 12822 | 8 | 396.0556 | 392.2783
3072 | 9646 | 4 | 595.9058 | 588.4551
4000 | 12610 | 4 | 779.0143 | 766.2403
4095 | 12916 | 4 | 797.9182 | 784.4403
HET 2 B2 BST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.82033 | =m | agrees with tag1_f.c
0 | - | - | -1.5425 | 0.57235 | =b
2 | 2067 | 512 | 0.99761 | 1.7403 | 0.74273
4 | 7330 | 512 | 3.5377 | 4.1784 | 0.64066
6 | 12714 | 512 | 6.1362 | 6.6164 | 0.48018
8 | 8938 | 256 | 8.6276 | 9.0545 | 0.42685
12 | 14041 | 256 | 13.5534 | 13.9306 | 0.37715
16 | 4780 | 64 | 18.456 | 18.8066 | 0.35061
24 | 7493 | 64 | 28.9312 | 28.5588 | -0.37237
32 | 10041 | 64 | 38.7692 | 38.311 | -0.45826
48 | 15148 | 64 | 58.4878 | 57.8153 | -0.67252
64 | 10128 | 32 | 78.2103 | 77.3197 | -0.89065
128 | 5075 | 8 | 156.7604 | 155.337 | -1.4234
256 | 10064 | 8 | 310.8644 | 311.3718 | 0.50739
512 | 10087 | 4 | 623.1496 | 623.4412 | 0.2916
1024 | 5077 | 1 | 1254.5775 | 1247.5802
2048 | 10260 | 1 | 2535.3486 | 2495.8581
3072 | 15474 | 1 | 3823.7801 | 3744.136
4095 | - | - | 5110.9534 | 4991.1949
HET 2 B2 BST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.82748 | =m | agrees with tag1_f.c
0 | - | - | -1.5406 | 0.33004 | =b
3 | 5607 | 512 | 2.7061 | 3.2266 | 0.52047
4 | 8540 | 512 | 4.1217 | 4.4351 | 0.31339
6 | 13566 | 512 | 6.5474 | 6.8521 | 0.30463
8 | 9351 | 256 | 9.0263 | 9.2691 | 0.24278
12 | 14529 | 256 | 14.0245 | 14.103 | 0.07855
16 | 4878 | 64 | 18.8344 | 18.937 | 0.10254
24 | 7394 | 64 | 28.5489 | 28.6049 | 0.055932
32 | 9946 | 64 | 38.4024 | 38.2728 | -0.12968
48 | 15016 | 64 | 57.9782 | 57.6086 | -0.36962
64 | 10022 | 32 | 77.3918 | 76.9444 | -0.4474
128 | 5008 | 8 | 154.6909 | 154.2876 | -0.40328
256 | 10032 | 8 | 309.8759 | 308.974 | -0.90194
512 | 9999 | 4 | 617.7132 | 618.3468 | 0.63362
1024 | 5012 | 1 | 1238.5153 | 1237.0925
2048 | 10113 | 1 | 2499.0234 | 2474.5839
2200 | 10871 | 1 | 2686.3328 | 2658.274
2400 | 11870 | 1 | 2933.1957 | 2899.9715
2600 | 12870 | 1 | 3180.3057 | 3141.6691
2800 | 13874 | 1 | 3428.4041 | 3383.3666
3072 | 15237 | 1 | 3765.2151 | 3712.0753
4095 | - | - | 5031.9709 | 4948.3581
HET 2 C1 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0674 | =m | agrees with tag1_f.c
0 | - | - | -1.4793 | 1.2589 | =b
2 | 1076 | 512 | 0.51932 | 0.69428 | 0.17496 | omitted pt at ch 1 with 0 MeV
4 | 5217 | 512 | 2.5179 | 2.5679 | 0.050007
6 | 8985 | 512 | 4.3365 | 4.4416 | 0.10508
8 | 12926 | 512 | 6.2386 | 6.3152 | 0.076654
12 | 10336 | 256 | 9.9771 | 10.0625 | 0.085444
16 | 14096 | 256 | 13.6065 | 13.8098 | 0.20331
24 | 5577 | 64 | 21.5333 | 21.3044 | -0.22893
32 | 7516 | 64 | 29.02 | 28.799 | -0.221
48 | 11398 | 64 | 44.0087 | 43.7882 | -0.22059
64 | 15279 | 64 | 58.9937 | 58.7773 | -0.21633
128 | 15388 | 32 | 118.829 | 118.734 | -0.094998
256 | 7713 | 8 | 238.2449 | 238.6474 | 0.40249
512 | 15494 | 8 | 478.5903 | 478.4742 | -0.11609
900 | 13685 | 4 | 845.4251 | 841.9617
1024 | 15600 | 4 | 963.729 | 958.1278
1200 | 4574 | 1 | 1130.2811 | 1123.0087
1400 | 5365 | 1 | 1325.7451 | 1310.3734
4095 | - | - | 3959.6227 | 3835.1124
HET 2 C1 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0878 | =m | did agree with tag1_f.c before mods
0 | - | - | -0.98747 | 0.63316 | =b
2 | - | - | 1.0111 | 1.2566 | 0.24543 | this pt scaled from HET 2 C1 AST HG room T assuming same slope | [5 empty] | 1.0111
3 | - | - | 2.0104 | 2.1759 | 0.16544 | this pt scaled from HET 2 C1 AST HG room T assuming same slope | [5 empty] | 2.0104
4 | - | - | 3.0097 | 3.0952 | 0.085453 | this pt scaled from HET 2 C1 AST HG room T assuming same slope | [5 empty] | 3.0097
6 | 10004 | 512 | 4.8283 | 4.9338 | 0.1055
8 | 13833 | 512 | 6.6763 | 6.7724 | 0.096112
12 | 11034 | 256 | 10.6508 | 10.4497 | -0.20115
16 | 14751 | 256 | 14.2387 | 14.1269 | -0.11182
24 | 5561 | 64 | 21.4715 | 21.4814 | 0.0098837
32 | 7503 | 64 | 28.9698 | 28.8359 | -0.13386
48 | 11342 | 64 | 43.7925 | 43.5449 | -0.2476
64 | 15131 | 64 | 58.4222 | 58.2539 | -0.16828
128 | 15173 | 32 | 117.1688 | 117.0899 | -0.07882
256 | 7590 | 8 | 234.4456 | 234.7619 | 0.31632
512 | 15222 | 8 | 470.1886 | 470.1059 | -0.082605
1024 | 15315 | 4 | 946.1224 | 940.794
1200 | 4488 | 1 | 1109.0297 | 1102.593
1400 | 5258 | 1 | 1299.3044 | 1286.4555
1500 | 5653 | 1 | 1396.9128 | 1378.3867
4095 | - | - | 3929.8521 | 3764.0028
HET 2 C1 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21407 | =m | agrees with tag1_f.c
0 | - | - | -7.7589 | 1.1409 | =b
2 | 2409 | 256 | 2.3253 | 4.0129 | 1.6875
4 | 3214 | 64 | 12.4096 | 13.3554 | 0.94586
6 | 5649 | 64 | 21.8113 | 22.698 | 0.88666
8 | 8138 | 64 | 31.4216 | 32.0405 | 0.61895
12 | 12972 | 64 | 50.0861 | 50.7256 | 0.63953
16 | 8870 | 32 | 68.4958 | 69.4107 | 0.91494
24 | 14057 | 32 | 108.5508 | 106.781 | -1.7698
32 | 4714 | 8 | 145.6096 | 144.1512 | -1.4584
48 | 7139 | 8 | 220.5148 | 218.8916 | -1.6232
64 | 9563 | 8 | 295.3891 | 293.6321 | -1.7571
128 | 9609 | 4 | 593.62 | 592.5938 | -1.0262
256 | 4808 | 1 | 1188.1049 | 1190.5173 | 2.4124
512 | 9659 | 1 | 2386.8355 | 2386.3643 | -0.47123
600 | 11336 | 1 | 2801.239 | 2797.4367
800 | 15165 | 1 | 3747.4232 | 3731.6921
4095 | - | - | 19335.8077 | 19123.5506
HET 2 C1 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21717 | =m | agrees with tag1_f.c
0 | - | - | -5.9306 | 0.73176 | =b
2 | 4612 | 256 | 4.4518 | 5.8399 | 1.3881
3 | 9990 | 256 | 9.6431 | 10.4447 | 0.80162
4 | 3922 | 64 | 15.1432 | 15.0495 | -0.093738
6 | 6122 | 64 | 23.6376 | 24.259 | 0.62139
8 | 8631 | 64 | 33.3251 | 33.4685 | 0.14345
12 | 13490 | 64 | 52.0862 | 51.8876 | -0.19853
16 | 9076 | 32 | 70.0866 | 70.3067 | 0.22013
24 | 13918 | 32 | 107.4774 | 107.1449 | -0.33254
32 | 4692 | 8 | 144.93 | 143.983 | -0.947
48 | 7091 | 8 | 219.0321 | 217.6593 | -1.3728
64 | 9472 | 8 | 292.5782 | 291.3356 | -1.2426
128 | 9498 | 4 | 586.7627 | 586.0409 | -0.7218
256 | 4747 | 1 | 1173.0312 | 1175.4514 | 2.4202
512 | 9530 | 1 | 2354.9583 | 2354.2724 | -0.68591
600 | 11180 | 1 | 2762.6898 | 2759.4921
800 | 14950 | 1 | 3694.2945 | 3680.446
4095 | - | - | 19042.4819 | 18853.1616
HET 2 C1 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 3.5955 | =m | agrees with tag1_f.c
0 | - | - | -0.54598 | 0.93627 | =b
4 | 6391 | 2048 | 0.77113 | 0.85211 | 0.080976 | probably off since phase = 0
5 | 9120 | 2048 | 1.1004 | 1.1302 | 0.029825 | probably off since phase = 0
6 | 11521 | 2048 | 1.3901 | 1.4084 | 0.01825 | probably off since phase = 0
8 | 4072 | 512 | 1.9653 | 1.9646 | -0.00067655 | probably off since phase = 0
12 | 6411 | 512 | 3.0942 | 3.0771 | -0.017053 | probably off since phase = 0
16 | 8704 | 512 | 4.2009 | 4.1896 | -0.011228 | probably off since phase = 0
24 | 13309 | 512 | 6.4234 | 6.4147 | -0.0087479
32 | 8984 | 256 | 8.672 | 8.6397 | -0.03233
48 | 13575 | 256 | 13.1036 | 13.0897 | -0.013857
64 | 4549 | 64 | 17.5641 | 17.5398 | -0.024341
128 | 9159 | 64 | 35.3638 | 35.34 | -0.023807
256 | 9189 | 32 | 70.9592 | 70.9403 | -0.018879
512 | 4601 | 8 | 142.1191 | 142.141 | 0.021868
1024 | 9245 | 8 | 285.5665 | 284.5424
2048 | 9286 | 4 | 573.6659 | 569.3452
3072 | 13991 | 4 | 864.329 | 854.148
3400 | 15505 | 4 | 957.8601 | 945.374
4095 | - | - | 1156.0435 | 1138.6727
HET 2 C1 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 3.6588 | =m | agrees with tag1_f.c
0 | - | - | -0.25266 | 0.52337 | =b
4 | 7240 | 2048 | 0.87357 | 0.9502 | 0.076628
6 | 11907 | 2048 | 1.4367 | 1.4968 | 0.060133
8 | 4185 | 512 | 2.0198 | 2.0434 | 0.023608
12 | 6483 | 512 | 3.1289 | 3.1367 | 0.0077524
16 | 8706 | 512 | 4.2018 | 4.2299 | 0.028094
24 | 13245 | 512 | 6.3925 | 6.4164 | 0.023892
32 | 8946 | 256 | 8.6353 | 8.6029 | -0.032434
48 | 13510 | 256 | 13.0408 | 12.9759 | -0.06497
64 | 4510 | 64 | 17.4135 | 17.3488 | -0.064686
128 | 9040 | 64 | 34.9043 | 34.8407 | -0.063551
256 | 9048 | 32 | 69.8704 | 69.8245 | -0.045836
512 | 4524 | 8 | 139.7407 | 139.7921 | 0.05137
1024 | 9083 | 8 | 280.5625 | 279.7272
1400 | 12443 | 8 | 384.3487 | 382.492
1800 | 16028 | 8 | 495.0849 | 491.8164
2048 | 9117 | 4 | 563.2255 | 559.5974
2200 | 9802 | 4 | 605.5431 | 601.1407
2700 | 12052 | 4 | 744.5424 | 737.7961
3072 | 13731 | 4 | 848.2669 | 839.4677
4095 | - | - | 1133.509 | 1119.0646
HET 2 C1 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.71634 | =m | agrees with tag1_f.c
0 | - | - | -2 | 0.80093 | =b
4 | 8594 | 512 | 4.1478 | 4.4659 | 0.31809 | probably off since phase = 0
6 | 14963 | 512 | 7.2217 | 7.2578 | 0.036152 | probably off since phase = 0
8 | 10406 | 256 | 10.0446 | 10.0498 | 0.0051858 | probably off since phase = 0
12 | 16269 | 256 | 15.704 | 15.6338 | -0.07026 | probably off since phase = 0
16 | 5520 | 64 | 21.3132 | 21.2177 | -0.095511 | probably off since phase = 0
24 | 8415 | 64 | 32.4911 | 32.3856 | -0.10547
32 | 11304 | 64 | 43.6458 | 43.5535 | -0.092266
48 | 8547 | 32 | 66.0015 | 65.8893 | -0.11219
64 | 11443 | 32 | 88.365 | 88.2252 | -0.13983
128 | 5743 | 8 | 177.3941 | 177.5684 | 0.17431
256 | 11528 | 8 | 356.0855 | 356.2549 | 0.16938
512 | 11553 | 4 | 713.7155 | 713.6279 | -0.087591
1024 | 5788 | 1 | 1430.2727 | 1428.3738
2048 | 11644 | 1 | 2877.3488 | 2857.8657
2800 | 15975 | 1 | 3947.5822 | 3907.6488
4095 | - | - | 5790.6039 | 5715.4535
HET 2 C1 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.72911 | =m | agrees with tag1_f.c
0 | - | - | -1.3929 | 0.49749 | =b
4 | 9262 | 512 | 4.4702 | 4.8038 | 0.33366
6 | 15336 | 512 | 7.4017 | 7.5469 | 0.1452
8 | 10537 | 256 | 10.1711 | 10.29 | 0.11892
11 | 14926 | 256 | 14.4077 | 14.4046 | -0.0030466
13 | 4453 | 64 | 17.1935 | 17.1477 | -0.045744
16 | 5492 | 64 | 21.2051 | 21.2623 | 0.057201
24 | 8363 | 64 | 32.2903 | 32.2347 | -0.055673
32 | 11258 | 64 | 43.4682 | 43.207 | -0.26121
64 | 11318 | 32 | 87.3997 | 87.0963 | -0.30344
128 | 5663 | 8 | 174.923 | 174.8749 | -0.048104
256 | 11345 | 8 | 350.4329 | 350.4321 | -0.00077046
512 | 11355 | 4 | 701.4835 | 701.5465 | 0.063007
1024 | 5686 | 1 | 1405.0675 | 1403.7754
2048 | 11427 | 1 | 2823.726 | 2808.233
2200 | 12286 | 1 | 3035.9935 | 3016.7072
2400 | 13414 | 1 | 3314.7335 | 3291.0154
2600 | 14543 | 1 | 3593.7207 | 3565.3235
2800 | 15673 | 1 | 3872.955 | 3839.6317
4095 | - | - | 5680.9971 | 5615.7769
HET 2 C2 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0834 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -1.4255 | 1.1624 | =b
2 | 4105 | 2048 | 0.49531 | 0.77312 | 0.27781
4 | 5006 | 512 | 2.4161 | 2.6192 | 0.20311
6 | 8913 | 512 | 4.3017 | 4.4653 | 0.16353
8 | 12771 | 512 | 6.1638 | 6.3113 | 0.14759
12 | 10333 | 256 | 9.9742 | 10.0035 | 0.029332
16 | 14116 | 256 | 13.6258 | 13.6957 | 0.069858
24 | 5545 | 64 | 21.4098 | 21.08 | -0.3298
32 | 7450 | 64 | 28.7651 | 28.4643 | -0.30087
48 | 11264 | 64 | 43.4914 | 43.2329 | -0.25846
64 | 15082 | 64 | 58.233 | 58.0015 | -0.2315
128 | 15173 | 32 | 117.1688 | 117.076 | -0.092751
256 | 7601 | 8 | 234.7854 | 235.225 | 0.43957
512 | 15269 | 8 | 471.6403 | 471.5229 | -0.11744
1000 | 14963 | 4 | 924.3767 | 921.9658
1500 | 5623 | 1 | 1389.4995 | 1383.4852
2000 | 7530 | 1 | 1860.7383 | 1845.0046
2200 | 8298 | 1 | 2050.5188 | 2029.6123
2400 | 9071 | 1 | 2241.5348 | 2214.2201
2600 | 9858 | 1 | 2436.0104 | 2398.8278
4095 | - | - | 3889.7153 | 3778.7708
HET 2 C2 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.1015 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -0.85994 | 0.74188 | =b
2 | 7491 | 2048 | 0.90386 | 1.1422 | 0.23835
3 | 3700 | 512 | 1.7858 | 2.0501 | 0.26432
4 | 6139 | 512 | 2.9629 | 2.9579 | -0.0049669
6 | 9633 | 512 | 4.6492 | 4.7737 | 0.12443
8 | 13472 | 512 | 6.5021 | 6.5894 | 0.087326
12 | 10726 | 256 | 10.3535 | 10.2209 | -0.13265
16 | 14358 | 256 | 13.8594 | 13.8523 | -0.0070529
24 | 5465 | 64 | 21.1009 | 21.1153 | 0.014401
32 | 7377 | 64 | 28.4833 | 28.3782 | -0.10507
48 | 11190 | 64 | 43.2056 | 42.9041 | -0.30155
64 | 14949 | 64 | 57.7195 | 57.43 | -0.28953
128 | 14979 | 32 | 115.6706 | 115.5334 | -0.13721
256 | 7493 | 8 | 231.4494 | 231.7404 | 0.291
512 | 15028 | 8 | 464.1961 | 464.1543 | -0.041793
575 | 8435 | 4 | 521.0932 | 521.35
1024 | 15083 | 4 | 931.79 | 928.9822
2048 | 7584 | 1 | 1874.0822 | 1858.638
2400 | 8916 | 1 | 2203.2328 | 2178.2071
4095 | - | - | 3788.2047 | 3778.7708
HET 2 C2 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21958 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -7.4326 | 1.1174 | =b
2 | 2226 | 256 | 2.1487 | 4.0195 | 1.8708
4 | 3038 | 64 | 11.73 | 13.1278 | 1.3978
6 | 5486 | 64 | 21.182 | 22.236 | 1.054
8 | 7877 | 64 | 30.4138 | 31.3442 | 0.93041
12 | 12755 | 64 | 49.2483 | 49.5607 | 0.31246
16 | 8731 | 32 | 67.4224 | 67.7772 | 0.35476
24 | 13730 | 32 | 106.0256 | 104.2101 | -1.8155
32 | 4603 | 8 | 142.1809 | 140.6431 | -1.5379
48 | 6962 | 8 | 215.0475 | 213.5089 | -1.5385
64 | 9323 | 8 | 287.9758 | 286.3748 | -1.601
128 | 9371 | 4 | 578.917 | 577.8383 | -1.0786
256 | 4690 | 1 | 1158.9459 | 1160.7654 | 1.8195
512 | 9416 | 1 | 2326.7878 | 2326.6195 | -0.16829
600 | 11049 | 1 | 2730.3184 | 2727.3818
800 | 14763 | 1 | 3648.0849 | 3638.2053
850 | 15696 | 1 | 3878.6386 | 3865.9112
4095 | - | - | 18841.5691 | 18644.0226
HET 2 C2 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.22291 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -5.727 | 0.74595 | =b
2 | 4337 | 256 | 4.1864 | 5.6258 | 1.4395
3 | 9472 | 256 | 9.1431 | 10.112 | 0.96894
4 | 3776 | 64 | 14.5795 | 14.5982 | 0.018675
6 | 5937 | 64 | 22.9233 | 23.5705 | 0.64717
8 | 8376 | 64 | 32.3405 | 32.5428 | 0.20228
12 | 13292 | 64 | 51.3217 | 50.4874 | -0.83422
16 | 8895 | 32 | 68.6889 | 68.4321 | -0.25678
24 | 13498 | 32 | 104.2341 | 104.3213 | 0.087258
32 | 4541 | 8 | 140.2658 | 140.2106 | -0.055196
48 | 6906 | 8 | 213.3177 | 211.9892 | -1.3285
64 | 9236 | 8 | 285.2885 | 283.7677 | -1.5208
128 | 9253 | 4 | 571.6272 | 570.8819 | -0.74532
256 | 4627 | 1 | 1143.378 | 1145.1102 | 1.7323
512 | 9283 | 1 | 2293.9221 | 2293.5669 | -0.35518
800 | 14547 | 1 | 3594.7092 | 3585.5807
4095 | - | - | 18476.9776 | 18367.4747
HET 2 C2 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0369 | =m | agrees with tag1_f.c
0 | - | - | -0.89457 | 0.69488 | =b
2 | 6689 | 2048 | 0.80709 | 1.2586 | 0.45154
4 | 5198 | 512 | 2.5087 | 3.1874 | 0.67863
6 | 10001 | 512 | 4.8268 | 5.1161 | 0.28928
8 | 14290 | 512 | 6.8969 | 7.0449 | 0.148
12 | 11212 | 256 | 10.8226 | 10.9024 | 0.079735
16 | 15324 | 256 | 14.7919 | 14.7599 | -0.031967
24 | 5930 | 64 | 22.8963 | 22.4749 | -0.4214
32 | 7940 | 64 | 30.6571 | 30.1899 | -0.46719
48 | 11932 | 64 | 46.0706 | 45.6199 | -0.45067
64 | 15912 | 64 | 61.4377 | 61.0499 | -0.38781
128 | 15925 | 32 | 122.9758 | 122.7699 | -0.20589
256 | 7962 | 8 | 245.9362 | 246.21 | 0.2738
512 | 15962 | 8 | 493.0462 | 493.0902 | 0.043948
1024 | 16023 | 4 | 989.8609 | 986.8505
1400 | 5481 | 1 | 1354.4099 | 1349.4557
1800 | 7072 | 1 | 1747.5619 | 1735.2059
2000 | 7872 | 1 | 1945.2499 | 1928.0811
2048 | 8064 | 1 | 1992.695 | 1974.3711
2200 | 8676 | 1 | 2143.9264 | 2120.9562
2400 | 9488 | 1 | 2344.5797 | 2313.8313 |  | omitted pt at ch 2400 not in tag1_f.c
4095 | - | - | 4045.1166 | 3948.4479
HET 2 C2 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0529 | =m | agrees with tag1_f.c
0 | - | - | 0.28608 | 0.22849 | =b
2 | 11825 | 2048 | 1.4268 | 1.6826 | 0.25577
3 | 4138 | 512 | 1.9972 | 2.6323 | 0.6352
4 | 6702 | 512 | 3.2346 | 3.5821 | 0.34751
6 | 11301 | 512 | 5.4543 | 5.4817 | 0.027435
8 | 15120 | 512 | 7.2975 | 7.3813 | 0.083821
12 | 11725 | 256 | 11.3178 | 11.1804 | -0.13739
16 | 15823 | 256 | 15.2735 | 14.9796 | -0.29393
24 | 5894 | 64 | 22.7573 | 22.5779 | -0.1794
32 | 7849 | 64 | 30.3057 | 30.1762 | -0.12953
48 | 11815 | 64 | 45.6188 | 45.3728 | -0.24603
64 | 15785 | 64 | 60.9474 | 60.5694 | -0.37797
128 | 15749 | 32 | 121.6167 | 121.3558 | -0.26093
256 | 7857 | 8 | 242.6929 | 242.9286 | 0.2357
512 | 15735 | 8 | 486.0345 | 486.0742 | 0.039758
1024 | 15785 | 4 | 975.1578 | 972.3655
1200 | 4619 | 1 | 1141.4011 | 1139.5281
1400 | 5397 | 1 | 1333.6527 | 1329.4856
1600 | 6177 | 1 | 1526.3985 | 1519.4431
1800 | 6961 | 1 | 1720.1327 | 1709.4007
2048 | 7937 | 1 | 1961.3121 | 1944.948
2200 | 8537 | 1 | 2109.5781 | 2089.3157
2400 | 9334 | 1 | 2306.5247 | 2279.2732
4095 | - | - | 3975.6478 | 3889.1632
HET 2 C2 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21001 | =m | agrees with tag1_f.c
0 | - | - | -5.4248 | 0.69535 | =b
2 | 3568 | 256 | 3.4441 | 6.2123 | 2.7682
4 | 3189 | 64 | 12.313 | 15.7357 | 3.4227
6 | 6308 | 64 | 24.3558 | 25.2591 | 0.90334
8 | 8803 | 64 | 33.9892 | 34.7825 | 0.7933
12 | 13868 | 64 | 53.5456 | 53.8293 | 0.28363
16 | 9469 | 32 | 73.1214 | 72.8761 | -0.24533
24 | 14646 | 32 | 113.0992 | 110.9696 | -2.1295
32 | 4895 | 8 | 151.2004 | 149.0632 | -2.1373
48 | 7361 | 8 | 227.3721 | 225.2503 | -2.1218
64 | 9825 | 8 | 303.482 | 301.4374 | -2.0446
128 | 9833 | 4 | 607.4582 | 606.1858 | -1.2724
256 | 4913 | 1 | 1214.0514 | 1215.6826 | 1.6312
512 | 9852 | 1 | 2434.5277 | 2434.6763 | 0.14854
800 | 15438 | 1 | 3814.8842 | 3806.0441
4095 | - | - | 19607.5041 | 19495.8257 |  | is 19605.7992 in tag1_f.c
HET 2 C2 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21313 | =m | agrees with tag1_f.c
0 | - | - | 1.6767 | 0.20172 | =b
2 | 7455 | 256 | 7.1961 | 8.4375 | 1.2413
3 | 10314 | 256 | 9.9558 | 13.1294 | 3.1736
4 | 3993 | 64 | 15.4173 | 17.8214 | 2.404
6 | 7103 | 64 | 27.4253 | 27.2053 | -0.22009
8 | 9408 | 64 | 36.3252 | 36.5892 | 0.264
12 | 14546 | 64 | 56.1635 | 55.357 | -0.80649
16 | 9807 | 32 | 75.7315 | 74.1248 | -1.6067
24 | 14590 | 32 | 112.6667 | 111.6604 | -1.0063
32 | 4847 | 8 | 149.7178 | 149.196 | -0.52174
48 | 7297 | 8 | 225.3952 | 224.2673 | -1.1279
64 | 9754 | 8 | 301.2889 | 299.3385 | -1.9503
128 | 9726 | 4 | 600.848 | 599.6235 | -1.2245
256 | 4852 | 1 | 1198.9777 | 1200.1934 | 1.2157
512 | 9717 | 1 | 2401.1679 | 2401.3333 | 0.16545
600 | 11394 | 1 | 2815.5713 | 2814.2252
700 | 13305 | 1 | 3287.7985 | 3283.4204
800 | 15218 | 1 | 3760.52 | 3752.6157
4095 | - | - | 19336.6911 | 19212.5998
HET 2 C3 AST HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0812 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -1.4665 | 1.232 | =b
2 | 4135 | 2048 | 0.49893 | 0.71029 | 0.21137
4 | 5106 | 512 | 2.4643 | 2.5601 | 0.095749
6 | 8870 | 512 | 4.281 | 4.4099 | 0.1289
8 | 12770 | 512 | 6.1633 | 6.2597 | 0.096417
12 | 10230 | 256 | 9.8747 | 9.9593 | 0.084536
16 | 13958 | 256 | 13.4733 | 13.6589 | 0.18559
24 | 5522 | 64 | 21.321 | 21.0581 | -0.26289
32 | 7433 | 64 | 28.6995 | 28.4573 | -0.24225
48 | 11264 | 64 | 43.4914 | 43.2556 | -0.23572
64 | 15097 | 64 | 58.2909 | 58.054 | -0.2369
128 | 15198 | 32 | 117.3618 | 117.2476 | -0.11425
256 | 7616 | 8 | 235.2487 | 235.6346 | 0.3859
512 | 15297 | 8 | 472.5052 | 472.4087 | -0.096462
1024 | 15371 | 4 | 949.582 | 945.957
2048 | 7742 | 1 | 1913.1256 | 1893.0535
2400 | 9112 | 1 | 2251.6663 | 2218.6179
2600 | 9907 | 1 | 2448.1188 | 2403.5977
4095 | - | - | 3916.6008 | 3786.3216
HET 2 C3 AST HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0982 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -0.90712 | 0.77415 | =b
2 | 7638 | 2048 | 0.92159 | 1.1162 | 0.19461
3 | 3804 | 512 | 1.836 | 2.0268 | 0.19081
4 | 6212 | 512 | 2.9981 | 2.9373 | -0.06083
6 | 9650 | 512 | 4.6574 | 4.7584 | 0.10097
8 | 13506 | 512 | 6.5185 | 6.5795 | 0.061032
12 | 10633 | 256 | 10.2638 | 10.2217 | -0.042014
16 | 14311 | 256 | 13.814 | 13.864 | 0.049925
24 | 5487 | 64 | 21.1858 | 21.1484 | -0.037439
32 | 7414 | 64 | 28.6261 | 28.4328 | -0.19334
48 | 11198 | 64 | 43.2365 | 43.0017 | -0.23486
64 | 14958 | 64 | 57.7542 | 57.5705 | -0.18371
128 | 15013 | 32 | 115.9332 | 115.846 | -0.087239
256 | 7513 | 8 | 232.0672 | 232.3968 | 0.32965
512 | 15073 | 8 | 465.5861 | 465.4986 | -0.087565
800 | 11798 | 4 | 728.8509 | 727.738
1024 | 15137 | 4 | 935.126 | 931.702
2048 | 7619 | 1 | 1882.7311 | 1864.109
2600 | 9735 | 1 | 2405.6159 | 2366.7346
2700 | 10130 | 1 | 2503.2243 | 2457.79
2800 | 12181 | 1 | 3010.0469 | 2548.8453
4095 | - | - | 9573.3997 | 3728.0123
HET 2 C3 AST LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21444 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -7.7183 | 1.2128 | =b
2 | 2354 | 256 | 2.2723 | 3.671 | 1.3987
4 | 3176 | 64 | 12.2628 | 12.9975 | 0.73469
6 | 5626 | 64 | 21.7225 | 22.3241 | 0.60157
8 | 8070 | 64 | 31.159 | 31.6507 | 0.49162
12 | 12917 | 64 | 49.8737 | 50.3038 | 0.43003
16 | 8824 | 32 | 68.1406 | 68.9569 | 0.81633
24 | 13954 | 32 | 107.7554 | 106.2632 | -1.4922
32 | 4687 | 8 | 144.7756 | 143.5694 | -1.2061
48 | 7104 | 8 | 219.4337 | 218.182 | -1.2517
64 | 9522 | 8 | 294.1227 | 292.7945 | -1.3282
128 | 9583 | 4 | 592.0138 | 591.2446 | -0.76919
256 | 4800 | 1 | 1186.128 | 1188.1448 | 2.0168
512 | 9641 | 1 | 2382.3875 | 2381.9452 | -0.4423
800 | 15118 | 1 | 3735.809 | 3724.9707
4095 | - | - | 19220.2664 | 19090.4877
HET 2 C3 AST LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.21782 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -5.9152 | 0.78453 | =b
2 | 4568 | 256 | 4.4094 | 5.5802 | 1.1708
3 | 9916 | 256 | 9.5717 | 10.1712 | 0.59953
4 | 3885 | 64 | 15.0003 | 14.7622 | -0.23819
6 | 6064 | 64 | 23.4137 | 23.9441 | 0.53043
8 | 8547 | 64 | 33.0008 | 33.1261 | 0.12529
12 | 13435 | 64 | 51.8738 | 51.49 | -0.38384
16 | 9018 | 32 | 69.6387 | 69.8538 | 0.21516
24 | 13819 | 32 | 106.7129 | 106.5816 | -0.13127
32 | 4659 | 8 | 143.9107 | 143.3094 | -0.60125
48 | 7058 | 8 | 218.0128 | 216.765 | -1.2478
64 | 9427 | 8 | 291.1882 | 290.2206 | -0.96764
128 | 9460 | 4 | 584.4152 | 584.043 | -0.3722
256 | 4734 | 1 | 1169.8187 | 1171.6876 | 1.8689
512 | 9500 | 1 | 2347.545 | 2346.977 | -0.56799
800 | 14890 | 1 | 3679.4679 | 3669.1776
4095 | - | - | 18917.9608 | 18796.4373
HET 2 C3 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0464 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -1.0239 | 0.76165 | =b
2 | 6373 | 2048 | 0.76896 | 1.1834 | 0.41447
4 | 5308 | 512 | 2.5618 | 3.0947 | 0.53289
6 | 10193 | 512 | 4.9195 | 5.006 | 0.086501
8 | 13942 | 512 | 6.7289 | 6.9173 | 0.18839
12 | 11141 | 256 | 10.7541 | 10.7399 | -0.014209
16 | 15001 | 256 | 14.4801 | 14.5625 | 0.082425
24 | 5855 | 64 | 22.6067 | 22.2077 | -0.39903
32 | 7828 | 64 | 30.2246 | 29.8529 | -0.37179
48 | 11780 | 64 | 45.4837 | 45.1432 | -0.34047
64 | 15729 | 64 | 60.7311 | 60.4336 | -0.29757
128 | 15769 | 32 | 121.7712 | 121.595 | -0.17617
256 | 7887 | 8 | 243.6196 | 243.9179 | 0.29831
512 | 15817 | 8 | 488.5674 | 488.5636 | -0.003732
1024 | 15884 | 4 | 981.2738 | 977.8551
1400 | 5435 | 1 | 1343.0429 | 1337.1786
1800 | 7014 | 1 | 1733.2295 | 1719.4375
2048 | 7998 | 1 | 1976.3858 | 1956.4381
2200 | 8608 | 1 | 2127.1229 | 2101.6965
2400 | 9417 | 1 | 2327.0349 | 2292.826
2600 | 11333 | 1 | 2800.4976 | 2483.9555
4095 | - | - | 6339.6318 | 3912.6484
HET 2 C3 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0608 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | 0.24892 | 0.27193 | =b
2 | 11725 | 2048 | 1.4147 | 1.629 | 0.21431
3 | 4139 | 512 | 1.9976 | 2.5717 | 0.57409
4 | 6613 | 512 | 3.1917 | 3.5144 | 0.32274
6 | 11311 | 512 | 5.4591 | 5.3998 | -0.059307
8 | 15022 | 512 | 7.2502 | 7.2852 | 0.035009
12 | 11619 | 256 | 11.2155 | 11.0559 | -0.15957
16 | 15541 | 256 | 15.0013 | 14.8267 | -0.17461
24 | 5815 | 64 | 22.4523 | 22.3682 | -0.084035
32 | 7779 | 64 | 30.0354 | 29.9097 | -0.1257
48 | 11739 | 64 | 45.3254 | 44.9928 | -0.33258
64 | 15644 | 64 | 60.403 | 60.0758 | -0.32711
128 | 15613 | 32 | 120.5665 | 120.408 | -0.15848
256 | 7795 | 8 | 240.7778 | 241.0724 | 0.2946
512 | 15618 | 8 | 482.4205 | 482.4012 | -0.019342
1024 | 15674 | 4 | 968.3005 | 965.0587
1200 | 4587 | 1 | 1133.4936 | 1130.9722
1400 | 5361 | 1 | 1324.7567 | 1319.5103
1600 | 6138 | 1 | 1516.7612 | 1508.0483
1800 | 6917 | 1 | 1709.2599 | 1696.5864
2048 | 7888 | 1 | 1949.2037 | 1930.3737
2100 | 8092 | 1 | 1999.6141 | 1979.3936
2300 | 8882 | 1 | 2194.831 | 2167.9316
2500 | 9681 | 1 | 2392.2719 | 2356.4697
2700 | 12089 | 1 | 2987.3128 | 2545.0078
4095 | - | - | 7137.7229 | 3860.061
HET 2 C3 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.20761 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -6.4461 | 0.77953 | =b
2 | 3275 | 256 | 3.1613 | 5.8787 | 2.7174
4 | 3307 | 64 | 12.7686 | 15.5121 | 2.7435
6 | 6542 | 64 | 25.2593 | 25.1456 | -0.11373
8 | 8780 | 64 | 33.9004 | 34.779 | 0.8786
12 | 14035 | 64 | 54.1905 | 54.0459 | -0.14455
16 | 9464 | 32 | 73.0828 | 73.3128 | 0.23001
24 | 14733 | 32 | 113.771 | 111.8466 | -1.9244
32 | 4923 | 8 | 152.0653 | 150.3804 | -1.6849
48 | 7414 | 8 | 229.0092 | 227.448 | -1.5612
64 | 9911 | 8 | 306.1384 | 304.5156 | -1.6228
128 | 9938 | 4 | 613.9448 | 612.7859 | -1.1589
256 | 4968 | 1 | 1227.6425 | 1229.3266 | 1.6842
512 | 9965 | 1 | 2462.4511 | 2462.4081 | -0.043073
800 | 15619 | 1 | 3859.6111 | 3849.6247
4095 | - | - | 19844.4792 | 19720.7315
HET 2 C3 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.2106 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | 1.6429 | 0.22148 | =b
2 | 7604 | 256 | 7.3399 | 8.4449 | 1.1049
3 | 10555 | 256 | 10.1885 | 13.1931 | 3.0047
4 | 4138 | 64 | 15.9772 | 17.9414 | 1.9642
6 | 7282 | 64 | 28.1165 | 27.4379 | -0.67855
8 | 9523 | 64 | 36.7692 | 36.9345 | 0.16527
12 | 14706 | 64 | 56.7812 | 55.9275 | -0.85371
16 | 9844 | 32 | 76.0172 | 74.9206 | -1.0966
24 | 14681 | 32 | 113.3694 | 112.9067 | -0.46271
32 | 4901 | 8 | 151.3858 | 150.8928 | -0.49292
48 | 7394 | 8 | 228.3914 | 226.8651 | -1.5263
64 | 9863 | 8 | 304.6557 | 302.8373 | -1.8184
128 | 9837 | 4 | 607.7053 | 606.7263 | -0.97892
256 | 4907 | 1 | 1212.5688 | 1214.5043 | 1.9356
512 | 9835 | 1 | 2430.3268 | 2430.0603 | -0.26652
600 | 11538 | 1 | 2851.1552 | 2847.9077
700 | 13474 | 1 | 3329.5601 | 3322.7343
800 | 15410 | 1 | 3807.9651 | 3797.5608
4095 | - | - | 19571.4085 | 19443.096
HET 2 C4 PEN HG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0303 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -1.1004 | 0.62444 | =b
2 | 5828 | 2048 | 0.7032 | 1.3351 | 0.63192
4 | 5194 | 512 | 2.5068 | 3.2763 | 0.7695
6 | 10043 | 512 | 4.8471 | 5.2175 | 0.3704
8 | 14350 | 512 | 6.9258 | 7.1587 | 0.23289
12 | 11306 | 256 | 10.9134 | 11.0411 | 0.12775
16 | 15459 | 256 | 14.9222 | 14.9235 | 0.0013715
24 | 5992 | 64 | 23.1357 | 22.6883 | -0.44733
32 | 8020 | 64 | 30.966 | 30.4532 | -0.51282
48 | 12048 | 64 | 46.5185 | 45.9828 | -0.53569
64 | 16068 | 64 | 62.0401 | 61.5124 | -0.52767
128 | 16065 | 32 | 124.0569 | 123.6309 | -0.42608
256 | 8021 | 8 | 247.7587 | 247.8678 | 0.10915
512 | 16062 | 8 | 496.1351 | 496.3417 | 0.20661
1024 | 16105 | 4 | 994.9266 | 993.2895
1600 | 6303 | 1 | 1557.5343 | 1552.3558
2048 | 8097 | 1 | 2000.8497 | 1987.1851
2200 | 8711 | 1 | 2152.5752 | 2134.7165
2400 | 9525 | 1 | 2353.7227 | 2328.8367
2600 | 10354 | 1 | 2558.5769 | 2522.9569
4095 | - | - | 4089.862 | 3974.0057
HET 2 C4 PEN HG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 1.0447 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | 0.28644 | 0.15627 | =b
2 | 11930 | 2048 | 1.4395 | 1.7648 | 0.32535
3 | 4177 | 512 | 2.016 | 2.722 | 0.70604
4 | 6487 | 512 | 3.1309 | 3.6792 | 0.54835
6 | 11377 | 512 | 5.491 | 5.5936 | 0.10266
8 | 15327 | 512 | 7.3974 | 7.508 | 0.11064
12 | 11841 | 256 | 11.4298 | 11.3368 | -0.092992
16 | 16016 | 256 | 15.4598 | 15.1656 | -0.29421
24 | 5968 | 64 | 23.043 | 22.8232 | -0.2198
32 | 7939 | 64 | 30.6532 | 30.4808 | -0.17242
48 | 11945 | 64 | 46.1208 | 45.796 | -0.32477
64 | 15956 | 64 | 61.6076 | 61.1112 | -0.49642
128 | 15906 | 32 | 122.8291 | 122.372 | -0.45715
256 | 7926 | 8 | 244.8242 | 244.8935 | 0.069272
512 | 15855 | 8 | 489.7411 | 489.9366 | 0.19546
1024 | 15889 | 4 | 981.5827 | 980.0228
1200 | 4648 | 1 | 1148.5673 | 1148.4899
1400 | 5429 | 1 | 1341.5602 | 1339.9298
1600 | 6211 | 1 | 1534.8002 | 1531.3697
1800 | 6997 | 1 | 1729.0287 | 1722.8096
2048 | 7974 | 1 | 1970.4551 | 1960.1951
2100 | 8178 | 1 | 2020.8656 | 2009.9695
2300 | 8972 | 1 | 2217.0709 | 2201.4094
2500 | 9772 | 1 | 2414.7589 | 2392.8493
4095 | - | - | 3991.3207 | 3919.5826
HET 2 C4 PEN LG
22 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.20599 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | -6.004 | 0.63787 | =b
2 | 3354 | 256 | 3.2375 | 6.6126 | 3.375
4 | 3232 | 64 | 12.4791 | 16.3218 | 3.8427
6 | 6429 | 64 | 24.823 | 26.0309 | 1.208
8 | 8984 | 64 | 34.6881 | 35.7401 | 1.0521
12 | 14160 | 64 | 54.6731 | 55.1585 | 0.48543
16 | 9647 | 32 | 74.4959 | 74.5769 | 0.080951
24 | 14959 | 32 | 115.5162 | 113.4137 | -2.1026
32 | 5004 | 8 | 154.5673 | 152.2504 | -2.3169
48 | 7529 | 8 | 232.5614 | 229.9239 | -2.6375
64 | 10045 | 8 | 310.2775 | 307.5974 | -2.68
128 | 10044 | 4 | 620.4932 | 618.2915 | -2.2017
256 | 5012 | 1 | 1238.5153 | 1239.6796 | 1.1643
512 | 10043 | 1 | 2481.7257 | 2482.4559 | 0.73018
800 | 15732 | 1 | 3887.5345 | 3880.5792
4095 | - | - | 19971.3538 | 19876.469
HET 2 C4 PEN LG
0 | C |  |  | ch = m*E + b
Chan | DAC | ATTN | MeV | 0.20942 | =m | not in tag1_f.c but agrees with fit in book
0 | - | - | 1.891 | 0.09203 | =b
2 | 7635 | 256 | 7.3699 | 9.1107 | 1.7409
3 | 10473 | 256 | 10.1093 | 13.8858 | 3.7765
4 | 4047 | 64 | 15.6258 | 18.6609 | 3.0351
6 | 7284 | 64 | 28.1242 | 28.2111 | 0.086883
8 | 9653 | 64 | 37.2711 | 37.7613 | 0.49013
12 | 14984 | 64 | 57.8546 | 56.8616 | -0.993
16 | 10035 | 32 | 77.4922 | 75.962 | -1.5302
24 | 14952 | 32 | 115.4621 | 114.1627 | -1.2994
32 | 4957 | 8 | 153.1155 | 152.3634 | -0.75209
48 | 7447 | 8 | 230.0285 | 228.7649 | -1.2636
64 | 9950 | 8 | 307.3431 | 305.1663 | -2.1767
128 | 9925 | 4 | 613.1417 | 610.7721 | -2.3696
256 | 4944 | 1 | 1221.7118 | 1221.9837 | 0.27189
512 | 9888 | 1 | 2443.4237 | 2444.4069 | 0.98323
600 | 11595 | 1 | 2865.2404 | 2864.6149
800 | 15479 | 1 | 3825.0157 | 3819.633
4095 | - | - | 19637.3128 | 19553.5564
Now format for tag1_f.c and tag1.c
float t2lal1 = 21.0 ; |  | /* see tag1_f.c for nomenclature */
int l2al1pt = 15 ; /* one less than the number of data points */
float l2al1chan[] = {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3500 | ,
4095
};
float l2al1mev[] = {
-0.20539 | ,
0.31906 | ,
0.49388 | ,
0.80171 | ,
1.1168 | ,
2.3801 | ,
3.6216 | ,
4.8736 | ,
9.8814 | ,
19.8595 | ,
39.8565 | ,
80.2286 | ,
161.3164 | ,
243.143 | ,
277.5439 | ,
325.3675
};
float t1lal1 = 0.0 ;
int l1al1pt = 16 ;
float l1al1chan[] = {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1al1mev[] = {
-0.15438 | ,
0.35286 | ,
0.52194 | ,
0.82896 | ,
1.1452 | ,
1.774 | ,
2.41 | ,
3.6721 | ,
4.9327 | ,
9.962 | ,
19.8832 | ,
39.8394 | ,
80.1083 | ,
160.9727 | ,
242.5244 | ,
316.378 | ,
324.2136
};
float t2lal2 = 21.0 ;
int l2al2pt = 19 ;
float l2al2chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
95 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l2al2mev[] = {
-0.083601 | ,
0.045072 | ,
0.10941 | ,
0.18636 | ,
0.34306 | ,
0.49281 | ,
0.78117 | ,
1.0758 | ,
2.2091 | ,
3.4004 | ,
4.5664 | ,
6.8003 | ,
9.2338 | ,
18.5665 | ,
37.2275 | ,
74.9018 | ,
150.4565 | ,
226.6127 | ,
295.3457 | ,
302.5627
};
float t1lal2 = 0.0 ;
int l1al2pt = 19 ;
float l1al2chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1al2mev[] = {
-0.063665 | ,
0.069371 | ,
0.13589 | ,
0.21523 | ,
0.37176 | ,
0.51791 | ,
0.8064 | ,
1.0981 | ,
1.6771 | ,
2.2612 | ,
3.4031 | ,
4.5503 | ,
9.1565 | ,
18.3625 | ,
36.7636 | ,
73.8966 | ,
148.3602 | ,
223.3479 | ,
291.0843 | ,
298.1638
};
float t2lal3 = 22.0 ;
int l2al3pt = 21 ;
float l2al3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2500 | ,
2600 | ,
3073 | ,
4000 | ,
4090 | ,
4095
};
float l2al3mev[] = {
-0.088805 | ,
0.80238 | ,
1.248 | ,
1.8446 | ,
3.1024 | ,
3.934 | ,
6.0132 | ,
8.0282 | ,
16.0815 | ,
24.5102 | ,
32.6842 | ,
65.322 | ,
130.4432 | ,
261.4733 | ,
524.1821 | ,
1050.4646 | ,
1284.4778 | ,
1335.8767 | ,
1582.2453 | ,
2067.0752 | ,
2114.026 | ,
2117.2385
};
float t1lal3 = 0.0 ;
int l1al3pt = 19 ;
float l1al3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1al3mev[] = {
0.1127 | ,
1.0039 | ,
1.4495 | ,
2.0333 | ,
3.2752 | ,
4.1087 | ,
6.2009 | ,
8.1903 | ,
12.1721 | ,
16.1432 | ,
24.2284 | ,
32.244 | ,
64.4494 | ,
128.3428 | ,
256.9017 | ,
514.483 | ,
1029.4603 | ,
1550.3681 | ,
2025.5607 | ,
2074.2413
};
float t2lbl1 = 22.0 ;
int l2bl1pt = 14 ;
float l2bl1chan[] = {
0 | ,
6 | ,
8 | ,
12 | ,
32 | ,
48 | ,
64 | ,
100 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
4000 | ,
4095
};
float l2bl1mev[] = {
-0.25063 | ,
0.29255 | ,
0.47361 | ,
0.7868 | ,
2.3616 | ,
3.6267 | ,
4.8865 | ,
7.7131 | ,
9.904 | ,
19.9025 | ,
39.8995 | ,
80.2716 | ,
161.2648 | ,
316.9966 | ,
324.7634
};
float t1lbl1 = 0.0 ;
int l1bl1pt = 16 ;
float l1bl1chan[] = {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1bl1mev[] = {
-0.17599 | ,
0.34041 | ,
0.51254 | ,
0.82117 | ,
1.134 | ,
1.7699 | ,
2.4046 | ,
3.6767 | ,
4.9402 | ,
9.9996 | ,
20.0077 | ,
40.0284 | ,
80.4004 | ,
161.4366 | ,
243.1086 | ,
316.9966 | ,
324.5606
};
float t2lbl2 = 22.0 ;
int l2bl2pt = 18 ;
float l2bl2chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l2bl2mev[] = {
-0.0885 | ,
0.042656 | ,
0.10823 | ,
0.18281 | ,
0.33585 | ,
0.48959 | ,
0.78318 | ,
1.0776 | ,
2.2644 | ,
3.4251 | ,
4.6062 | ,
9.3133 | ,
18.7297 | ,
37.5712 | ,
75.5891 | ,
151.8656 | ,
228.6919 | ,
298.1638 | ,
305.3807
};
float t1lbl2 = 0.0 ;
int l1bl2pt = 19 ;
float l1bl2chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1bl2mev[] = {
-0.059806 | ,
0.068867 | ,
0.1332 | ,
0.21378 | ,
0.36991 | ,
0.51818 | ,
0.81083 | ,
1.1038 | ,
1.6944 | ,
2.287 | ,
3.445 | ,
4.6051 | ,
9.2414 | ,
18.5364 | ,
37.1244 | ,
74.6183 | ,
149.8208 | ,
225.5302 | ,
293.9023 | ,
301.0505
};
float t2lbl3 = 22.0 ;
int l2bl3pt = 18 ;
float l2bl3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l2bl3mev[] = {
0.057434 | ,
0.94862 | ,
1.3942 | ,
1.9909 | ,
3.2486 | ,
4.1019 | ,
6.297 | ,
8.4297 | ,
16.9425 | ,
25.8269 | ,
34.4294 | ,
68.8742 | ,
137.6403 | ,
276.0219 | ,
553.65 | ,
1109.5239 | ,
1671.9463 | ,
2185.4408 | ,
2238.5695
};
float t1lbl3 = 0.0 ;
int l1bl3pt = 19 ;
float l1bl3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1bl3mev[] = {
0.26135 | ,
1.1525 | ,
1.5981 | ,
2.182 | ,
3.4682 | ,
4.3408 | ,
6.5827 | ,
8.7261 | ,
12.8825 | ,
17.1046 | ,
25.6222 | ,
34.1938 | ,
68.017 | ,
135.4781 | ,
271.1106 | ,
542.9624 | ,
1088.0253 | ,
1637.598 | ,
2139.7255 | ,
2191.3715
};
float t2lcl1 = 22.0 ;
int l2cl1pt = 16 ;
float l2cl1chan[] = {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3500 | ,
4095
};
float l2cl1mev[] = {
-0.1574 | ,
0.18297 | ,
0.35316 | ,
0.51926 | ,
0.79956 | ,
1.1118 | ,
2.2776 | ,
3.4294 | ,
4.5836 | ,
9.1941 | ,
18.3216 | ,
36.5831 | ,
73.4326 | ,
147.3464 | ,
221.6639 | ,
252.9031 | ,
296.3313
};
float t1lcl1 = 0.0 ;
int l1cl1pt = 17 ;
float l1cl1chan[] = {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1cl1mev[] = {
-0.10293 | ,
0.23744 | ,
0.40763 | ,
0.55819 | ,
0.83688 | ,
1.1328 | ,
1.7109 | ,
2.3012 | ,
3.4689 | ,
4.6341 | ,
9.2478 | ,
18.4097 | ,
36.7722 | ,
73.7419 | ,
147.8962 | ,
222.4887 | ,
289.8471 | ,
296.8578
};
float t2lcl2 = 22.0 ;
int l2cl2pt = 18 ;
float l2cl2chan[] = {
0 | ,
1 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l2cl2mev[] = {
-0.039199 | ,
0.01923 | ,
0.07766 | ,
0.12928 | ,
0.21429 | ,
0.37478 | ,
0.53416 | ,
1.1423 | ,
2.3925 | ,
3.6104 | ,
4.8285 | ,
9.6946 | ,
19.43 | ,
38.8771 | ,
78.0463 | ,
156.6081 | ,
235.8401 | ,
307.374 | ,
314.8659
};
float t1lcl2 = 0.0 ;
int l1cl2pt = 20 ;
float l1cl2chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2500 | ,
3072 | ,
4000 | ,
4095
};
float l1cl2mev[] = {
-0.064538 | ,
0.083668 | ,
0.15777 | ,
0.24624 | ,
0.40589 | ,
0.55738 | ,
0.86829 | ,
1.1705 | ,
1.7793 | ,
2.3895 | ,
3.5787 | ,
4.7694 | ,
9.6097 | ,
19.2023 | ,
38.3874 | ,
77.0153 | ,
154.4946 | ,
188.8955 | ,
232.5581 | ,
302.9751 | ,
310.3295
};
float t2lcl3 = 22.0 ;
int l2cl3pt = 22 ;
float l2cl3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2200 | ,
2400 | ,
2600 | ,
3072 | ,
4000 | ,
4090 | ,
4095
};
float l2cl3mev[] = {
-1.338 | ,
0.11619 | ,
0.84329 | ,
1.5097 | ,
2.6477 | ,
3.7873 | ,
8.2647 | ,
13.1132 | ,
17.5178 | ,
26.4292 | ,
35.329 | ,
70.8897 | ,
141.8103 | ,
284.4545 | ,
570.8859 | ,
1145.602 | ,
1231.8434 | ,
1344.7726 | ,
1458.6903 | ,
1727.0518 | ,
2259.3267 | ,
2311.7141 | ,
2314.6794
};
float t1lcl3 = 0.0 ;
int l1cl3pt = 19 ;
float l1cl3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1cl3mev[] = {
-0.9958 | ,
0.45838 | ,
1.1855 | ,
1.8519 | ,
3.2226 | ,
4.1763 | ,
6.4196 | ,
8.5784 | ,
12.8825 | ,
17.2243 | ,
26.0624 | ,
34.9004 | ,
69.7777 | ,
139.37 | ,
279.2343 | ,
559.9513 | ,
1123.3621 | ,
1692.4564 | ,
2212.6229 | ,
2265.9987
};
float t2ldl1 = 22.0 ;
int l2dl1pt = 18 ;
float l2dl1chan[] = {
0 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
728 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4090 | ,
4095
};
float l2dl1mev[] = {
-0.16291 | ,
0.17747 | ,
0.34766 | ,
0.49214 | ,
1.0846 | ,
1.7433 | ,
2.3385 | ,
3.5075 | ,
4.6771 | ,
9.3724 | ,
18.7663 | ,
37.5454 | ,
53.5 | ,
75.4431 | ,
151.5219 | ,
228.1076 | ,
297.4077 | ,
304.2123 | ,
304.5559
};
float t1ldl1 = 0.0 ;
int l1dl1pt = 17 ;
float l1dl1chan[] = {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1dl1mev[] = {
-0.1032 | ,
0.23718 | ,
0.40736 | ,
0.5508 | ,
0.84668 | ,
1.1408 | ,
1.7222 | ,
2.3055 | ,
3.4466 | ,
4.603 | ,
9.2618 | ,
18.5386 | ,
37.0729 | ,
74.455 | ,
149.4771 | ,
224.9803 | ,
293.215 | ,
300.2945
};
float t2ldl2 = 22.0 ;
int l2dl2pt = 19 ;
float l2dl2chan[] = {
0 | ,
1 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l2dl2mev[] = {
-0.053664 | ,
0.0047657 | ,
0.063195 | ,
0.11481 | ,
0.19982 | ,
0.36031 | ,
0.51832 | ,
1.1521 | ,
1.8519 | ,
2.4865 | ,
3.7422 | ,
4.9982 | ,
10.0286 | ,
20.0765 | ,
40.1229 | ,
80.4778 | ,
161.3335 | ,
242.8852 | ,
316.4467 | ,
324.1448
};
float t1ldl2 = 0.0 ;
int l1dl2pt = 19 ;
float l1dl2chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l1dl2mev[] = {
-0.061148 | ,
0.085111 | ,
0.15824 | ,
0.25295 | ,
0.41213 | ,
0.57201 | ,
0.89084 | ,
1.2042 | ,
1.8319 | ,
2.4636 | ,
3.692 | ,
4.9251 | ,
9.8922 | ,
19.808 | ,
39.5644 | ,
79.3265 | ,
159.031 | ,
239.2252 | ,
311.6354 | ,
319.1961
};
float t2ldl3 = 22.0 ;
int l2dl3pt = 18 ;
float l2dl3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
};
float l2dl3mev[] = {
-0.94259 | ,
0.28283 | ,
0.89553 | ,
1.3963 | ,
2.599 | ,
3.6511 | ,
7.8573 | ,
12.3449 | ,
16.5332 | ,
24.8114 | ,
33.051 | ,
66.0788 | ,
132.0494 | ,
264.6857 | ,
530.7923 | ,
1064.5499 | ,
1603.7439 | ,
2097.9639 | ,
2148.6215
};
float t1ldl3 = 0.0 ;
int l1dl3pt = 21 ;
float l1dl3chan[] = {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2500 | ,
3072 | ,
3500 | ,
4000 | ,
4095
};
float l1dl3mev[] = {
-0.68197 | ,
0.54345 | ,
1.1562 | ,
1.6569 | ,
2.9504 | ,
3.9943 | ,
6.0499 | ,
8.1421 | ,
12.2966 | ,
16.429 | ,
24.5643 | ,
32.5954 | ,
65.0672 | ,
130.1343 | ,
260.6702 | ,
522.3288 | ,
1047.4993 | ,
1280.524 | ,
1577.3031 | ,
1799.9492 | ,
2061.6387 | ,
2111.8021
};
float | t2ha1aH | = | 22 | ;
int | h2a1aHpt | = | 15 | ;
float | h2a1aHchan[] | = | {
0 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2a1aHmev[] | = | {
-0.054402 | ,
0.32517 | ,
0.51496 | ,
0.69834 | ,
1.068 | ,
1.4536 | ,
2.2438 | ,
2.951 | ,
5.9551 | ,
11.923 | ,
23.9406 | ,
48.01 | ,
96.6472 | ,
145.2844 | ,
189.8405 | ,
194.4285
} | ;
float | t1ha1aH | = | 0 | ;
int | h1a1aHpt | = | 16 | ;
float | h1a1aHchan[] | = | {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h1a1aHmev[] | = | {
-0.032185 | ,
0.25355 | ,
0.3488 | ,
0.53859 | ,
0.72196 | ,
1.0917 | ,
1.4606 | ,
2.2134 | ,
2.9561 | ,
5.8874 | ,
11.7737 | ,
23.6098 | ,
47.3313 | ,
95.1265 | ,
143.0162 | ,
186.8163 | ,
191.3527
} | ;
float | t2ha1aL | = | 22 | ;
int | h2a1aLpt | = | 16 | ;
float | h2a1aLchan[] | = | {
0 | ,
4 | ,
5 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2a1aLmev[] | = | {
-0.080815 | ,
0.8113 | ,
1.0343 | ,
1.7034 | ,
3.4877 | ,
5.271 | ,
7.0945 | ,
10.6697 | ,
14.2449 | ,
28.6423 | ,
57.3835 | ,
115.2566 | ,
231.3209 | ,
463.7416 | ,
698.7397 | ,
913.0492 | ,
935.1812
} | ;
float | t1ha1aL | = | 0 | ;
int | h1a1aLpt | = | 15 | ;
float | h1a1aLchan[] | = | {
0 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h1a1aLmev[] | = | {
-0.12632 | ,
1.6723 | ,
2.5716 | ,
3.4565 | ,
5.2398 | ,
7.0161 | ,
10.5666 | ,
14.087 | ,
28.1784 | ,
56.5071 | ,
113.4095 | ,
227.5234 | ,
456.3871 | ,
687.1238 | ,
897.653 | ,
919.5101
} | ;
float | t2ha2aH | = | 22 | ;
int h2a2aHpt | = | 18 | ;
float | h2a2aHchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2a2aHmev[] | = | {
-0.059906 | ,
0.027386 | ,
0.11468 | ,
0.20982 | ,
0.30336 | ,
0.49966 | ,
0.68652 | ,
1.067 | ,
1.4383 | ,
2.1957 | ,
2.9418 | ,
5.9261 | ,
11.8811 | ,
23.8461 | ,
47.8639 | ,
96.3379 | ,
144.9923 | ,
189.5828 | ,
194.2395
} | ;
float | t1ha2aH | = | 0 | ;
int | h1a2aHpt | = | 17 | ;
float | h1a2aHchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h1a2aHmev[] | = | {
-0.057591 | ,
0.15082 | ,
0.25503 | ,
0.34494 | ,
0.53845 | ,
0.7221 | ,
1.095 | ,
1.4665 | ,
2.2156 | ,
2.9483 | ,
5.9035 | ,
11.8221 | ,
23.6978 | ,
47.5289 | ,
95.1952 | ,
145.5937 | ,
188.2253 | ,
192.7789
} | ;
float | t2ha2aL | = | 22 | ;
int | h2a2aLpt | = | 17 | ;
float | h2a2aLchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
200 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2a2aLmev[] | = | {
-0.061887 | ,
0.69095 | ,
1.0674 | ,
1.5182 | ,
3.344 | ,
5.243 | ,
7.0322 | ,
10.6128 | ,
14.1934 | ,
28.5435 | ,
44.6163 | ,
57.2202 | ,
114.9474 | ,
230.9085 | ,
463.7416 | ,
699.0833 | ,
914.1489 | ,
936.3497
} | ;
float | t1ha2aL | = | 0 | ;
int | h1a2aLpt | = | 19 | ;
float | h1a2aLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h1a2aLmev[] | = | {
-0.16965 | ,
0.24953 | ,
0.45911 | ,
0.71579 | ,
1.2154 | ,
1.6214 | ,
2.5122 | ,
3.3929 | ,
5.1646 | ,
6.9227 | ,
10.4324 | ,
13.9421 | ,
28.0452 | ,
56.215 | ,
112.8682 | ,
226.6471 | ,
450.2699 | ,
682.5874 | ,
893.4603 | ,
915.7298
} | ;
float | t2hb1pH | = | 22 | ;
int | h2b1pHpt | = | 17 | ;
float | h2b1pHchan[] | = | {
0 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2b1pHmev[] | = | {
-0.42243 | ,
0.12923 | ,
0.31311 | ,
0.70525 | ,
1.1092 | ,
2.6448 | ,
4.1299 | ,
5.6377 | ,
8.7155 | ,
11.7628 | ,
23.9581 | ,
48.3872 | ,
97.3845 | ,
195.3713 | ,
393.8007 | ,
593.1258 | ,
775.9254 | ,
794.8293
} | ;
float | t1hb1pH | = | 0 | ;
int | h1b1pHpt | = | 24 | ;
float | h1b1pHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1000 | ,
1024 | ,
1536 | ,
2048 | ,
2500 | ,
3072 | ,
3500 | ,
4000 | ,
4090 | ,
4095
} | ;
float | h1b1pHmev[] | = | {
-0.24301 | ,
0.077222 | ,
0.23734 | ,
0.43051 | ,
0.84848 | ,
1.2486 | ,
1.9856 | ,
2.7491 | ,
4.2559 | ,
5.7414 | ,
8.7145 | ,
11.6837 | ,
23.5913 | ,
47.5648 | ,
95.5157 | ,
187.2476 | ,
191.7574 | ,
288.6863 | ,
385.9858 | ,
472.1654 | ,
580.8321 | ,
662.9344 | ,
759.5544 | ,
777.0992 | ,
777.9023
} | ;
float | t2hb1pL | = | 22 | ;
int | h2b1pLpt | = | 17 | ;
float | h2b1pLchan[] | = | {
0 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2b1pLmev[] | = | {
-1.516 | ,
0.41845 | ,
1.0632 | ,
2.4103 | ,
3.7877 | ,
8.8805 | ,
13.8999 | ,
19.0391 | ,
29.271 | ,
39.4951 | ,
80.3494 | ,
161.9497 | ,
326.1234 | ,
654.7179 | ,
1316.1079 | ,
1985.5288 | ,
2596.879 | ,
2660.6334
} | ;
float | t1hb1pL | = | 0 | ;
int | h1b1pLpt | = | 19 | ;
float | h1b1pLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h1b1pLmev[] | = | {
-0.76981 | ,
0.25869 | ,
0.77294 | ,
1.4711 | ,
2.8891 | ,
4.1965 | ,
6.6609 | ,
9.2357 | ,
14.2503 | ,
19.2514 | ,
29.2439 | ,
39.1785 | ,
79.2142 | ,
159.2006 | ,
320.0075 | ,
642.7331 | ,
1290.6555 | ,
1945.0028 | ,
2543.009 | ,
2605.0336
} | ;
float | t2hb2bH | = | 22 | ;
int | h2b2bHpt | = | 18 | ;
float | h2b2bHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
10 | ,
16 | ,
64 | ,
100 | ,
128 | ,
256 | ,
512 | ,
1000 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h2b2bHmev[] | = | {
-0.17194 | ,
0.14105 | ,
0.29755 | ,
0.48541 | ,
0.87478 | ,
1.2121 | ,
1.5343 | ,
2.6724 | ,
11.8487 | ,
17.8923 | ,
23.4407 | ,
46.6652 | ,
93.2609 | ,
182.2127 | ,
186.8152 | ,
375.7616 | ,
565.5112 | ,
739.4767 | ,
757.5157
} | ;
float | t1hb2bH | = | 0 | ;
int | h1b2bHpt | = | 20 | ;
float | h1b2bHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
100 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | h1b2bHmev[] | = | {
-0.012669 | ,
0.2745 | ,
0.41808 | ,
0.62743 | ,
1.0144 | ,
1.323 | ,
2.0724 | ,
2.7737 | ,
4.1825 | ,
5.5875 | ,
8.4519 | ,
11.2754 | ,
17.6877 | ,
22.6646 | ,
45.4258 | ,
91.0832 | ,
182.7687 | ,
367.7923 | ,
553.5264 | ,
723.5381 | ,
741.0829
} | ;
float | t2hb2bL | = | 22 | ;
int | h2b2bLpt | = | 18 | ;
float | h2b2bLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3200 | ,
4095
} | ;
float | h2b2bLmev[] | = | {
-0.71672 | ,
1.0565 | ,
1.9431 | ,
3.3331 | ,
5.9369 | ,
8.0813 | ,
17.3788 | ,
27.7227 | ,
37.1167 | ,
55.9395 | ,
74.7585 | ,
149.934 | ,
300.8564 | ,
603.4426 | ,
1210.839 | ,
2437.9873 | ,
3675.5141 | ,
3830.9463 | ,
4917.7573
} | ;
float | t1hb2bL | = | 0 | ;
int | h1b2bLpt | = | 15 | ;
float | h1b2bLchan[] | = | {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4095
} | ;
float | h1b2bLmev[] | = | {
0.17954 | ,
6.5084 | ,
8.618 | ,
13.2339 | ,
17.8691 | ,
27.0894 | ,
36.2943 | ,
54.7078 | ,
73.1754 | ,
146.9378 | ,
294.9258 | ,
591.2725 | ,
1185.8809 | ,
2387.5768 | ,
3597.6745 | ,
4806.5904
} | ;
float | t2hc1pH | = | 22 | ;
int | h2c1pHpt | = | 16 | ;
float | h2c1pHchan[] | = | {
0 | ,
5 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3500 | ,
4090 | ,
4095
} | ;
float | h2c1pHmev[] | = | {
-0.56167 | ,
0.80359 | ,
1.0766 | ,
1.6106 | ,
2.6424 | ,
3.6849 | ,
7.8892 | ,
16.1664 | ,
32.7228 | ,
65.8162 | ,
131.9259 | ,
264.0679 | ,
530.5452 | ,
799.7715 | ,
913.1332 | ,
1067.5152 | ,
1068.9979
} | ;
float | t1hc1pH | = | 0 | ;
int | h1c1pHpt | = | 18 | ;
float | h1c1pHchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3250 | ,
3800 | ,
4095
} | ;
float | h1c1pHmev[] | = | {
-0.31456 | ,
0.74435 | ,
1.2738 | ,
1.808 | ,
2.8196 | ,
3.8317 | ,
5.8689 | ,
7.9288 | ,
11.979 | ,
16.0313 | ,
32.1668 | ,
64.5498 | ,
129.3003 | ,
259.0639 | ,
519.8577 | ,
782.9063 | ,
828.8687 | ,
971.513 | ,
1048.0222
} | ;
float | t2hc1pL | = | 22 | ;
int | h2c1pLpt | = | 10 | ;
float | h2c1pLchan[] | = | {
0 | ,
5 | ,
6 | ,
8 | ,
12 | ,
16 | ,
256 | ,
1024 | ,
2048 | ,
3072 | ,
4095
} | ;
float | h2c1pLmev[] | = | {
-3.0059 | ,
4.2723 | ,
5.7279 | ,
8.1701 | ,
13.2802 | ,
18.5062 | ,
326.2779 | ,
1305.4821 | ,
2629.2504 | ,
3963.1502 | ,
5295.7473
} | ;
float | t1hc1pL | = | 0 | ;
int | h1c1pLpt | = | 16 | ;
float | h1c1pLchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4095
} | ;
float | h1c1pLmev[] | = | {
-1.4942 | ,
3.7955 | ,
6.4403 | ,
8.9519 | ,
14.0283 | ,
19.1433 | ,
29.298 | ,
39.4372 | ,
59.5381 | ,
79.4922 | ,
159.4477 | ,
320.1001 | ,
641.4358 | ,
1281.0182 | ,
2575.8746 | ,
3879.627 | ,
5182.1062
} | ;
float | t2Ha1aH | = | 22 | ;
int | H2a1aHpt | = | 17 | ;
float | H2a1aHchan[] | = | {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3500 | ,
4000 | ,
4090 | ,
4095
} | ;
float | H2a1aHmev[] | = | {
-0.10307 | ,
0.20905 | ,
0.31309 | ,
0.5163 | ,
0.71847 | ,
1.5325 | ,
2.313 | ,
3.0949 | ,
6.2321 | ,
12.5223 | ,
25.1434 | ,
50.493 | ,
101.6819 | ,
153.0684 | ,
174.7193 | ,
200.2364 | ,
204.8759 | ,
205.1335
} | ;
float | t1Ha1aH | = | 0 | ;
int | H1a1aHpt | = | 16 | ;
float | H1a1aHchan[] | = | {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1a1aHmev[] | = | {
-0.069606 | ,
0.23758 | ,
0.33997 | ,
0.54181 | ,
0.73539 | ,
1.1247 | ,
1.517 | ,
2.2958 | ,
3.0648 | ,
6.1602 | ,
12.3601 | ,
24.789 | ,
49.7455 | ,
100.1268 | ,
150.6971 | ,
197.0403 | ,
201.8516
} | ;
float | t2Ha1aL | = | 22 | ;
int | H2a1aLpt | = | 18 | ;
float | H2a1aLchan[] | = | {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3500 | ,
4000 | ,
4090 | ,
4095
} | ;
float | H2a1aLmev[] | = | {
-0.27654 | ,
1.1105 | ,
1.5728 | ,
2.5756 | ,
3.5201 | ,
5.5792 | ,
7.4973 | ,
11.3474 | ,
15.1997 | ,
30.6764 | ,
61.5762 | ,
123.6936 | ,
248.4698 | ,
499.2078 | ,
752.5576 | ,
859.0938 | ,
984.4628 | ,
1007.3508 | ,
1008.5193
} | ;
float | t1Ha1aL | = | 0 | ;
int | H1a1aLpt | = | 16 | ;
float | H1a1aLchan[] | = | {
0 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1a1aLmev[] | = | {
-0.060276 | ,
1.2277 | ,
1.657 | ,
2.6615 | ,
3.5902 | ,
5.5105 | ,
7.4532 | ,
11.2711 | ,
15.0547 | ,
30.2661 | ,
60.6569 | ,
121.709 | ,
244.2943 | ,
490.4787 | ,
739.0859 | ,
966.3172 | ,
989.8239
} | ;
float | t2Ha2aH | = | 22 | ;
int | H2a2aHpt | = | 20 | ;
float | H2a2aHchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
100 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3500 | ,
4000 | ,
4095
} | ;
float | H2a2aHmev[] | = | {
-0.064269 | ,
0.025708 | ,
0.11568 | ,
0.22241 | ,
0.3245 | ,
0.52879 | ,
0.73176 | ,
1.1423 | ,
1.5343 | ,
2.3388 | ,
3.1343 | ,
4.908 | ,
6.2869 | ,
12.5932 | ,
25.2529 | ,
50.639 | ,
101.8279 | ,
153.1715 | ,
174.7708 | ,
200.3051 | ,
205.1852
} | ;
float | t1Ha2aH | = | 0 | ;
int | H1a2aHpt | = | 19 | ;
float | H1a2aHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1a2aHmev[] | = | {
-0.029836 | ,
0.053798 | ,
0.095615 | ,
0.14663 | ,
0.25292 | ,
0.34692 | ,
0.55537 | ,
0.7519 | ,
1.1385 | ,
1.5309 | ,
2.3256 | ,
3.1112 | ,
6.2246 | ,
12.4428 | ,
24.9179 | ,
49.9173 | ,
100.3244 | ,
150.8174 | ,
197.1606 | ,
201.9547
} | ;
float | t2Ha2aL | = | 22 | ;
int | H2a2aLpt | = | 18 | ;
float | H2a2aLchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H2a2aLmev[] | = | {
-0.31809 | ,
0.13636 | ,
0.59081 | ,
1.0898 | ,
1.5521 | ,
2.5222 | ,
3.464 | ,
5.4922 | ,
7.3834 | ,
11.1133 | ,
14.8624 | ,
29.7958 | ,
59.6345 | ,
119.6126 | ,
239.9641 | ,
481.4747 | ,
725.0644 | ,
948.3092 | ,
971.6097
} | ;
float | t1Ha2aL | = | 0 | ;
int | H1a2aLpt | = | 19 | ;
float | H1a2aLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1a2aLmev[] | = | {
-0.26809 | ,
0.21627 | ,
0.45844 | ,
0.72371 | ,
1.1718 | ,
1.6525 | ,
2.6178 | ,
3.5336 | ,
5.3665 | ,
7.2116 | ,
10.935 | ,
14.6358 | ,
29.3168 | ,
58.6378 | ,
117.5162 | ,
235.5995 | ,
472.5394 | ,
711.3178 | ,
929.8888 | ,
952.5706
} | ;
float | t2Hb1pH | = | 22 | ;
int | H2b1pHpt | = | 17 | ;
float | H2b1pHchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H2b1pHmev[] | = | {
-0.38309 | ,
0.46973 | ,
0.89614 | ,
1.2926 | ,
2.056 | ,
2.822 | ,
4.4321 | ,
5.9345 | ,
8.9741 | ,
11.9684 | ,
23.8809 | ,
47.7772 | ,
95.786 | ,
192.1589 | ,
386.9125 | ,
582.3147 | ,
761.5312 | ,
780.0645
} | ;
float | t1Hb1pH | = | 0 | ;
int | H1b1pHpt | = | 19 | ;
float | H1b1pHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1b1pHmev[] | = | {
-0.056107 | ,
0.23999 | ,
0.38804 | ,
0.58073 | ,
0.99508 | ,
1.3656 | ,
2.1448 | ,
2.8944 | ,
4.3365 | ,
5.8076 | ,
8.7917 | ,
11.7541 | ,
23.4793 | ,
46.9625 | ,
94.1103 | ,
188.8229 | ,
379.9316 | ,
571.3801 | ,
746.9518 | ,
765.2379
} | ;
float | t2Hb1pL | = | 22 | ;
int | H2b1pLpt | = | 17 | ;
float | H2b1pLchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H2b1pLmev[] | = | {
-1.2756 | ,
1.4899 | ,
2.8727 | ,
4.0913 | ,
6.4924 | ,
8.9491 | ,
14.0882 | ,
18.9425 | ,
28.5026 | ,
38.0086 | ,
75.9245 | ,
151.8182 | ,
304.7175 | ,
611.9061 | ,
1229.3722 | ,
1852.3366 | ,
2422.6664 | ,
2481.9728
} | ;
float | t1Hb1pL | = | 0 | ;
int | H1b1pLpt | = | 19 | ;
float | H1b1pLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1b1pLmev[] | = | {
-0.19402 | ,
0.74978 | ,
1.2217 | ,
1.8832 | ,
3.2467 | ,
4.4383 | ,
6.8192 | ,
9.2425 | ,
13.8835 | ,
18.5912 | ,
28.074 | ,
37.5105 | ,
74.8589 | ,
149.5016 | ,
299.7444 | ,
601.5275 | ,
1207.3795 | ,
1819.4709 | ,
2377.9395 | ,
2435.7633
} | ;
float | t2Hb2bH | = | 22 | ;
int | H2b2bHpt | = | 18 | ;
float | H2b2bHchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H2b2bHmev[] | = | {
-0.07324 | ,
0.26521 | ,
0.60366 | ,
1.017 | ,
1.445 | ,
2.2558 | ,
3.0363 | ,
4.656 | ,
6.2072 | ,
9.3294 | ,
12.454 | ,
24.9156 | ,
49.4567 | ,
98.7359 | ,
198.9544 | ,
401.8935 | ,
605.1106 | ,
791.1844 | ,
810.459
} | ;
float | t1Hb2bH | = | 0 | ;
int | H1b2bHpt | = | 18 | ;
float | H1b2bHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
2048 | ,
3072 | ,
4000 | ,
4095
} | ;
float | H1b2bHmev[] | = | {
-0.091701 | ,
0.25725 | ,
0.43172 | ,
0.64167 | ,
1.0512 | ,
1.4092 | ,
2.2356 | ,
2.9933 | ,
4.5238 | ,
6.0725 | ,
9.1798 | ,
12.2358 | ,
24.5257 | ,
49.0668 | ,
97.9328 | ,
396.0556 | ,
595.9058 | ,
779.0143 | ,
797.9182
} | ;
float | t2Hb2bL | = | 22 | ;
int | H2b2bLpt | = | 17 | ;
float | H2b2bLchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
4095
} | ;
float | H2b2bLmev[] | = | {
-1.5425 | ,
0.99761 | ,
3.5377 | ,
6.1362 | ,
8.6276 | ,
13.5534 | ,
18.456 | ,
28.9312 | ,
38.7692 | ,
58.4878 | ,
78.2103 | ,
156.7604 | ,
310.8644 | ,
623.1496 | ,
1254.5775 | ,
2535.3486 | ,
3823.7801 | ,
5110.9534
} | ;
float | t1Hb2bL | = | 0 | ;
int | H1b2bLpt | = | 21 | ;
float | H1b2bLchan[] | = | {
0 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2200 | ,
2400 | ,
2600 | ,
2800 | ,
3072 | ,
4095
} | ;
float | H1b2bLmev[] | = | {
-1.5406 | ,
2.7061 | ,
4.1217 | ,
6.5474 | ,
9.0263 | ,
14.0245 | ,
18.8344 | ,
28.5489 | ,
38.4024 | ,
57.9782 | ,
77.3918 | ,
154.6909 | ,
309.8759 | ,
617.7132 | ,
1238.5153 | ,
2499.0234 | ,
2686.3328 | ,
2933.1957 | ,
3180.3057 | ,
3428.4041 | ,
3765.2151 | ,
5031.9709
} | ;
float | t2Hc1pH | = | 22 | ;
int | H2c1pHpt | = | 18 | ;
float | H2c1pHchan[] | = | {
0 | ,
4 | ,
5 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
3072 | ,
3400 | ,
4095
} | ;
float | H2c1pHmev[] | = | {
-0.54598 | ,
0.77113 | ,
1.1004 | ,
1.3901 | ,
1.9653 | ,
3.0942 | ,
4.2009 | ,
6.4234 | ,
8.672 | ,
13.1036 | ,
17.5641 | ,
35.3638 | ,
70.9592 | ,
142.1191 | ,
285.5665 | ,
573.6659 | ,
864.329 | ,
957.8601 | ,
1156.0435
} | ;
float | t1Hc1pH | = | 0 | ;
int | H1c1pHpt | = | 20 | ;
float | H1c1pHchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
1400 | ,
1800 | ,
2048 | ,
2200 | ,
2700 | ,
3072 | ,
4095
} | ;
float | H1c1pHmev[] | = | {
-0.25266 | ,
0.87357 | ,
1.4367 | ,
2.0198 | ,
3.1289 | ,
4.2018 | ,
6.3925 | ,
8.6353 | ,
13.0408 | ,
17.4135 | ,
34.9043 | ,
69.8704 | ,
139.7407 | ,
280.5625 | ,
384.3487 | ,
495.0849 | ,
563.2255 | ,
605.5431 | ,
744.5424 | ,
848.2669 | ,
1133.509
} | ;
float | t2Hc1pL | = | 22 | ;
int | H2c1pLpt | = | 16 | ;
float | H2c1pLchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2800 | ,
4095
} | ;
float | H2c1pLmev[] | = | {
-2 | ,
4.1478 | ,
7.2217 | ,
10.0446 | ,
15.704 | ,
21.3132 | ,
32.4911 | ,
43.6458 | ,
66.0015 | ,
88.365 | ,
177.3941 | ,
356.0855 | ,
713.7155 | ,
1430.2727 | ,
2877.3488 | ,
3947.5822 | ,
5790.6039
} | ;
float | t1Hc1pL | = | 0 | ;
int | H1c1pLpt | = | 19 | ;
float | H1c1pLchan[] | = | {
0 | ,
4 | ,
6 | ,
8 | ,
11 | ,
13 | ,
16 | ,
24 | ,
32 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2200 | ,
2400 | ,
2600 | ,
2800 | ,
4095
} | ;
float | H1c1pLmev[] | = | {
-1.3929 | ,
4.4702 | ,
7.4017 | ,
10.1711 | ,
14.4077 | ,
17.1935 | ,
21.2051 | ,
32.2903 | ,
43.4682 | ,
87.3997 | ,
174.923 | ,
350.4329 | ,
701.4835 | ,
1405.0675 | ,
2823.726 | ,
3035.9935 | ,
3314.7335 | ,
3593.7207 | ,
3872.955 | ,
5680.9971
} | ;
float | t2hcaH | = | 22 | ; | /*HET 1 C1 AST HG 22C */
int | h2caHpt | = | 17 | ;
float | h2caHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
32 | ,
64 | ,
100 | ,
200 | ,
400 | ,
600 | ,
800 | ,
1000 | ,
1200 | ,
1400 | ,
4095
} | ;
float | h2caHmev[] | = | {
-0.94657 | ,
0.21538 | ,
0.79635 | ,
1.4764 | ,
3.5643 | ,
5.8515 | ,
13.425 | ,
28.8385 | ,
58.6539 | ,
92.033 | ,
183.8498 | ,
368.5646 | ,
554.3295 | ,
741.9478 | ,
931.0487 | ,
1119.1612 | ,
1318.579 | ,
4005.7334
} | ;
float | t1hcaH | = | 0 | ;
int | h1caHpt | = | 18 | ;
float | h1caHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
1200 | ,
1400 | ,
4095
} | ;
float | h1caHmev[] | = | {
-0.36886 | ,
0.68788 | ,
1.2162 | ,
1.9416 | ,
3.7568 | ,
6.2159 | ,
9.8854 | ,
13.6326 | ,
21.0545 | ,
28.4331 | ,
43.0821 | ,
57.5535 | ,
115.3231 | ,
230.9243 | ,
463.393 | ,
933.2727 | ,
1094.6973 | ,
1283.9836 | ,
3834.6159
} | ;
float | t2hcaL | = | 22 | ;
int | h2caLpt | = | 13 | ;
float | h2caLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
32 | ,
64 | ,
128 | ,
256 | ,
512 | ,
800 | ,
4095
} | ;
float | h2caLmev[] | = | {
-5.81 | ,
1.113 | ,
4.5744 | ,
8.5176 | ,
19.5642 | ,
29.1744 | ,
66.8278 | ,
142.8914 | ,
290.7867 | ,
584.4152 | ,
1163.641 | ,
2341.6144 | ,
3678.7266 | ,
18976.5902
} | ;
float | t1hcaL | = | 0 | ;
int | h1caLpt | = | 16 | ;
float | h1caLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
800 | ,
4095
} | ;
float | h1caLmev[] | = | {
-4.1594 | ,
2.92 | ,
6.4596 | ,
10.8574 | ,
21.4252 | ,
31.3135 | ,
49.1556 | ,
67.6464 | ,
103.9947 | ,
140.2658 | ,
212.8235 | ,
284.7016 | ,
571.0094 | ,
1140.1655 | ,
2291.9453 | ,
3597.6745 | ,
18536.4865
} | ;
float | t2hcpH | = | 22 | ; | /* HET 1 C2 PEN HG 22C */
int | h2cpHpt | = | 19 | ;
float | h2cpHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
2200 | ,
2400 | ,
2600 | ,
4095
} | ;
float | h2cpHmev[] | = | {
-0.22298 | ,
1.1535 | ,
1.8417 | ,
2.9171 | ,
5.0436 | ,
6.5547 | ,
13.8671 | ,
21.6221 | ,
28.854 | ,
43.3215 | ,
57.7929 | ,
115.5703 | ,
230.8316 | ,
461.5706 | ,
922.5852 | ,
1858.7614 | ,
1999.8612 | ,
2186.1822 | ,
2374.48 | ,
3782.0062
} | ;
float | t1hcpH | = | 0 | ;
int | h1cpHpt | = | 17 | ;
float | h1cpHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
2048 | ,
4095
} | ;
float | h1cpHmev[] | = | {
-0.24433 | ,
1.4601 | ,
2.3123 | ,
3.4639 | ,
5.4982 | ,
6.9601 | ,
10.7377 | ,
14.2764 | ,
21.4677 | ,
28.5721 | ,
42.6844 | ,
56.8121 | ,
113.6938 | ,
226.9396 | ,
453.6322 | ,
906.0288 | ,
1826.1429 | ,
3665.4725
} | ;
float | t2hcpL | = | 22 | ;
int | h2cpLpt | = | 7 | ;
float | h2cpLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
16 | ,
4095
} | ;
float | h2cpLmev[] | = | {
-4.2723 | ,
4.9383 | ,
9.5437 | ,
14.3633 | ,
25.1898 | ,
33.4023 | ,
70.2642 | ,
19113.7766
} | ;
float | t1hcpL | = | 0 | ;
int | h1cpLpt | = | 16 | ;
float | h1cpLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
800 | ,
4095
} | ;
float | h1cpLmev[] | = | {
-0.76643 | ,
7.3998 | ,
11.4829 | ,
17.3904 | ,
27.8346 | ,
35.3638 | ,
54.1364 | ,
71.994 | ,
108.2573 | ,
143.9107 | ,
215.0784 | ,
286.3696 | ,
572.5539 | ,
1142.1424 | ,
2279.0955 | ,
3563.3262 | ,
18256.1736
} | ;
float | t2HcaH | = | 22 | ; | /* HET 2 C1 AST HG 22C */
int | H2caHpt | = | 18 | ;
float | H2caHchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
900 | ,
1024 | ,
1200 | ,
1400 | ,
4095
} | ;
float | H2caHmev[] | = | {
-1.4793 | ,
0.51932 | ,
2.5179 | ,
4.3365 | ,
6.2386 | ,
9.9771 | ,
13.6065 | ,
21.5333 | ,
29.02 | ,
44.0087 | ,
58.9937 | ,
118.829 | ,
238.2449 | ,
478.5903 | ,
845.4251 | ,
963.729 | ,
1130.2811 | ,
1325.7451 | ,
3959.6227
} | ;
float | t1HcaH | = | 0 | ;
int | H1caHpt | = | 19 | ;
float | H1caHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
1200 | ,
1400 | ,
1500 | ,
4095
} | ;
float | H1caHmev[] | = | {
-0.98747 | ,
1.0111 | ,
2.0104 | ,
3.0097 | ,
4.8283 | ,
6.6763 | ,
10.6508 | ,
14.2387 | ,
21.4715 | ,
28.9698 | ,
43.7925 | ,
58.4222 | ,
117.1688 | ,
234.4456 | ,
470.1886 | ,
946.1224 | ,
1109.0297 | ,
1299.3044 | ,
1396.9128 | ,
3929.8521
} | ;
float | t2HcaL | = | 22 | ;
int | H2caLpt | = | 16 | ;
float | H2caLchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
600 | ,
800 | ,
4095
} | ;
float | H2caLmev[] | = | {
-7.7589 | ,
2.3253 | ,
12.4096 | ,
21.8113 | ,
31.4216 | ,
50.0861 | ,
68.4958 | ,
108.5508 | ,
145.6096 | ,
220.5148 | ,
295.3891 | ,
593.62 | ,
1188.1049 | ,
2386.8355 | ,
2801.239 | ,
3747.4232 | ,
19335.8077
} | ;
float | t1HcaL | = | 0 | ;
int | H1caLpt | = | 17 | ;
float | H1caLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
600 | ,
800 | ,
4095
} | ;
float | H1caLmev[] | = | {
-5.9306 | ,
4.4518 | ,
9.6431 | ,
15.1432 | ,
23.6376 | ,
33.3251 | ,
52.0862 | ,
70.0866 | ,
107.4774 | ,
144.93 | ,
219.0321 | ,
292.5782 | ,
586.7627 | ,
1173.0312 | ,
2354.9583 | ,
2762.6898 | ,
3694.2945 | ,
19042.4819
} | ;
float | t2HcpH | = | 22 | ;
int | H2cpHpt | = | 21 | ;
float | H2cpHchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
1400 | ,
1800 | ,
2000 | ,
2048 | ,
2200 | ,
2400 | ,
4095
} | ;
float | H2cpHmev[] | = | {
-0.89457 | ,
0.80709 | ,
2.5087 | ,
4.8268 | ,
6.8969 | ,
10.8226 | ,
14.7919 | ,
22.8963 | ,
30.6571 | ,
46.0706 | ,
61.4377 | ,
122.9758 | ,
245.9362 | ,
493.0462 | ,
989.8609 | ,
1354.4099 | ,
1747.5619 | ,
1945.2499 | ,
1992.695 | ,
2143.9264 | ,
2344.5797 | ,
4045.1166
} | ;
float | t1HcpH | = | 0 | ;
int | H1cpHpt | = | 23 | ;
float | H1cpHchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
1024 | ,
1200 | ,
1400 | ,
1600 | ,
1800 | ,
2048 | ,
2200 | ,
2400 | ,
4095
} | ;
float | H1cpHmev[] | = | {
0.28608 | ,
1.4268 | ,
1.9972 | ,
3.2346 | ,
5.4543 | ,
7.2975 | ,
11.3178 | ,
15.2735 | ,
22.7573 | ,
30.3057 | ,
45.6188 | ,
60.9474 | ,
121.6167 | ,
242.6929 | ,
486.0345 | ,
975.1578 | ,
1141.4011 | ,
1333.6527 | ,
1526.3985 | ,
1720.1327 | ,
1961.3121 | ,
2109.5781 | ,
2306.5247 | ,
3975.6478
} | ;
float | t2HcpL | = | 22 | ;
int | H2cpLpt | = | 15 | ;
float | H2cpLchan[] | = | {
0 | ,
2 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
800 | ,
4095
} | ;
float | H2cpLmev[] | = | {
-5.4248 | ,
3.4441 | ,
12.313 | ,
24.3558 | ,
33.9892 | ,
53.5456 | ,
73.1214 | ,
113.0992 | ,
151.2004 | ,
227.3721 | ,
303.482 | ,
607.4582 | ,
1214.0514 | ,
2434.5277 | ,
3814.8842 | ,
19607.5041
} | ;
int | H1cpLpt | = | 18 | ;
float | t1HcpL | = | 0 | ;
float | H1cpLchan[] | = | {
0 | ,
2 | ,
3 | ,
4 | ,
6 | ,
8 | ,
12 | ,
16 | ,
24 | ,
32 | ,
48 | ,
64 | ,
128 | ,
256 | ,
512 | ,
600 | ,
700 | ,
800 | ,
4095
} | ;
float | H1cpLmev[] | = | {
1.6767 | ,
7.1961 | ,
9.9558 | ,
15.4173 | ,
27.4253 | ,
36.3252 | ,
56.1635 | ,
75.7315 | ,
112.6667 | ,
149.7178 | ,
225.3952 | ,
301.2889 | ,
600.848 | ,
1198.9777 | ,
2401.1679 | ,
2815.5713 | ,
3287.7985 | ,
3760.52 | ,
19336.6911
} | ;
float | t1td1 | = | 0 | ;
float | t1td2 | = | 0 | ;
float | t2td1 | = | 22 | ;
float | t2td2 | = | 22 | ;